Echo Planar Imaging Pulse Sequences

There are several enhancements to the Diffusion Weighted Echo Planar

Imaging (DW EPI) sequence. There is also a new addition to the Echo

Planar pulse sequence family, Diffusion Tensor Imaging (DTI).

Diffusion Weighted EPI

Increased gradient strength enables DW EPI sensitivities to extend beyond

the standard B-value of 1000 s/mm2. The maximum B-value is now 7000

s/mm2. This allows greater diffusion weighting, increasing the sensitivity of

detection of non-mobile protons.

The Dual Spin Echo feature was moved from the User CV window to the

Diffusion Options window. You can perform a dual-spin echo DW EPI

acquisition by changing the User Control Variable (CV) 12, Dual Spin Echo

Dif., from 0 (off) to 1 (on). When on, an additional refocusing pulse is

applied in the diffusion sequence before the readout. This additional pulse

reduces the eddy currents and minimizes image distortion; however, it

increases the minimum TE by 10 to 14 ms. This option only produces a

single echo.

Diffusion Tensor EPI

Diffusion Tensor Imaging (DTI) is the newest addition to the Echo Planar

pulse sequence family. It is used to visualize the white matter fiber tracts in

the central nervous system (CNS) and provides information on molecular

diffusion of biological tissues.

Refer to the Echo Planar Imaging Pulse Sequences chapter in this guide for

additional information on DTI.

Fast Gradient Echo

With the advent of faster hardware, electrocardiogram (ECG) gating, and

K-space segmenting, Fast GRE sequences have enabled the division of

K-space into segments and its acquisition in successive cardiac cycles,

thereby effectively "freezing" cardiac motion. Typically, these are

two-dimensional (2D) sequences and are within the realm of

breath-holding, in order to reduce respiratory motion artifacts.

Coronary artery imaging involves imaging of tortuous vessels with

diameters ranging from 2 to 4 mm. This imposes an even higher spatial

resolution limit on the imaging sequences. Three-dimensional (3D) imaging

sequences are advantageous, as they allow high-resolution data sets that

permit reformatting of the slices post acquisition. This can help visualize the

length of tortuous vessels, which may be missed with a conventional 2D

technique. 3D imaging times, however, are in the order of minutes, which

rule out breath-holding as a viable option. Early sequences used respiratory

bellows to track and gate to the motion of the diaphragm. This technique is

sub-optimal as the abdominal motion is not necessarily representative of

the diaphragmatic motion, especially in the instance where someone is a

"chest breather." The technique of choice is to use navigator echoes to

directly estimate the position of the right hemi-diaphragm and gate the

acquisition based on this. A Fast GRE sequence with perpendicularly

intersecting slab planes is used to excite a column of spins. Fourier

transform of the resultant echo yields the position of the liver-lung edge,

enabling prospective gating from the diaphragm position.

In prospective gating, the current position of the right hemi-diaphragm (as

collected from the phase of the current navigator echo in relation to a

tracker reference) is compared against the threshold as deduced from a

histogram of the diaphragm displacement data over a period of time. If the

position is within a narrow acceptance window about the threshold, the

acquired data from the main sequence (3D Fast GRE) following the

navigator sequence is accepted. The acceptance window size is fixed to an

optimal value to provide a trade-off between the total scan time and the

suppression of motion artifacts. Reducing the window size generally results

in better image quality. However, if the window size is reduced too far,

image degradation may result, directly as a product of increased scan times

and changes in the respiratory cycle. Based on the total scan time, it is

likely the respiratory waveform may undergo a shift in its baseline. If this

occurs, the scan may not complete in an acceptable time, unless the

acceptance window is widened and/or the threshold is shifted accordingly.

The Navigator option allows you to manually adjust the acceptance window

and/or the threshold value during the scan acquisition.

# Cine IR

Determine optimal TI time for your myocardial delayed enhancement study.

## Cine IR

Cine IR captures image contrast evolution at different TI times to quickly determine optimal TI value for myocardial delayed enhancement (MDE) in a rapid, single breath-hold 20 second acquisition. Multiple TI evolution images can also help you detect certain tissue T1 abnormalities such as amyloid, vasculitis or diffuse fibrosis.

It uses multiphase FGRE Cine acquisition that captures image contrast evolution at different TI times. It also features Adiabatic Inversion Recovery (IR) for uniform myocardial suppression, and supports both 1 RR and 2 RR mode Multiphase FGRE Cine acquisition.

Fast Spin Echo & Flair

It is possible to prescribe multiple groups of slice locations, each with a

unique tilt angle within the same acquisition, on all GE Signa® MRI

systems. This capability is referred to as Multi-slice, Multi-angle, or MSMA.

MSMA is available for any 2D Fast Spin Echo (FSE), Spin Echo (SE), and

multi-planar Gradient Echo (GRE) sequences. This is useful when an entire

range of slice locations is not desired or needed for a given anatomical

body part (e.g., axial spine and sagittal TMJ examinations).

Note that dark banding artifacts, commonly known as cross-talk, and low

SNR tissue contrast changes are likely to occur when MSMA slices

intersect. The data for multiple groups are generally acquired in an

interleaved slice acquisition order. For example, two groups of five slice

locations are prescribed and the slices intersect one another. The first

group contains image numbers 1 to 5; the second group, 6 to 10.

Interleaved slice acquisition order results in the image data being gathered

as slices 1/3/5/7 and 9 gathered first, then the system goes back and

collects data for 2/4/6/8/ and 10. Collecting the data in this manner results in

cross-talk where the groups intersect (Figure 3-8). Care should be taken to

avoid overlapping groups on top of the anatomy of interest.

NOTE: Sequential GRE and sequential Fast GRE/SPGR sequences do not

exhibit this cross-talk because the slices are acquired one at a time.

Time of Flight

Time-of-Flight (ToF) Imaging refers to the process of measuring the depth of a

scene by quantifying the changes that an emitted light signal encounters when it

bounces back from objects in a scene.

Phase Contrast Vascular

Phase-contrast MRI (PC-MRI), a noninvasive technique with which flow can be measured accurately with flexible spatial and temporal resolution [1, 2], has a variety of established applications in quantifying cardiovascular function and hemodynamics. Current and emerging cardiovascular applications are reviewed.

Research

Proactive Service

Research PSD

Spectroscopy/PROBE

Spectroscopy Prescan

The Spectroscopy Prescan window, Figure 6-9, consists of two display

windows, each with a display control panel just below the window, and a

region at the left of the screen that allows for parameter entry, adjustment,

and control. A typical PROSE spectrum from the entire PROSE volume is

shown in both display windows in Figure 6-9.

You can visualize and examine the spectrum acquired from the entire

PROSE volume in the Spectroscopy Prescan window. It may be possible to

improve the homogeneity through the PROSE volume by manually

adjusting the X, Y, and Z gradient currents in the Gradient Shimming area

on the Spectroscopy Prescan window.

Videoconferencing

SGD_EchoSpeed

Diffusion Weighted EPI

Increased gradient strength enables DW EPI sensitivities to extend beyond

the standard B-value of 1000 s/mm2. The maximum B-value is now 7000

s/mm2. This allows greater diffusion weighting, increasing the sensitivity of

detection of non-mobile protons.

The Dual Spin Echo feature was moved from the User CV window to the

Diffusion Options window. You can perform a dual-spin echo DW EPI

acquisition by changing the User Control Variable (CV) 12, Dual Spin Echo

Dif., from 0 (off) to 1 (on). When on, an additional refocusing pulse is

applied in the diffusion sequence before the readout. This additional pulse

reduces the eddy currents and minimizes image distortion; however, it

increases the minimum TE by 10 to 14 ms. This option only produces a

single echo.

Flair EPI

flair but no flair epi

SPECIAL

SPECIAL is Spectral Inversion at Lipids is a fat suppression technique that

uses a 180° inversion pulse to nullify the signal from fat. This option is not

compatible with Elliptic-centric or Reverse Elliptic-centric k-space filling.

SPECIAL is turned on when Fat SAT is selected from the Saturation

window. Selection of SPECIAL results in a Prep Time display on the Scan

Timing screen.

Smart Prep

SmartStep (Multi Station) is an optional feature that provides automatic

table movement and switching of prescribed coils between stations,

including body to phased array, and/or between phased array elements.

SmartStep is used with 3D Fast Time Of Flight Gradient Echo (TOF

GRE)/SPoiled Gradient Echo (TOF SPGR), and 3D Fast GRE/SPGR pulse

sequences.

SSFSE

Single Shot Fast Spin Echo-Excel (SSFSE-XL) is a purchasable option

available for your MR system. It is a modification of the original SSFSE

pulse sequence. This feature is designed to improve the CNR and SNR

seen in MRCP images.

The SSFSE-XL sequence is activated by selecting SSFSE from the Pulse

Sequence window. If you have this feature, the original SSFSE pulse

sequence is not available to you.

Three Plane Localizer

Localizer Viewports

When you open 3-Plane GRx, a busy cursor appears on the screen until

initialization is complete and the appropriate three localizer images display.

The localizer images can be images from a 3-Plane Localizer or, if another

series is used, three images of the same plane display. A prescription can

be made on any of the three localizer images. The intersections of the

slices prescribed are shown on the other viewports. For example, if a

graphic prescription is made on the Viewport1 localizer image, the slice

intersections are shown on Viewport2 and Viewport3 localizer images. The

intersections signify the area on the localizer image that will be scanned.

There are various scenarios 3-Plane GRx uses for determining which

localizer images are displayed as default images in the three viewports. The

default load is the last valid 3-Plane Localizer.

• If you are opening 3-Plane GRx for the first time, the middle image of

the coronal set of the most recent series of the 3-Plane Localizer loads

in the first viewport. The middle image of the sagittal set loads in

Viewport2 and the middle image of the axial set loads in Viewport3.

– If the localizer series is not a 3-Plane Localizer, the last compatible

series scanned is divided into thirds and loads in the viewports,

displaying the center slices of each third.

• If you are entering 3-Plane GRx for the second time and have acquired

a 3-Plane Localizer, the system displays this series.

– If this localizer series is not a 3-Plane Localizer, the valid localized

series for the set prescription displays.

• If you had previously prescribed SAT bands, then the default images are

the localizers that were used in the SAT prescription.

At any instant of time one of the three viewports is always highlighted. The

active viewport is highlighted with a green or bright border. There are

various graphic tools that can be applied either to a selected viewport or to

all the viewports. The following tools operate on the images loaded into the

viewports:

• Window and Level (W/L) – using the center mouse button

• Pan – using the right mouse button

• Zoom – using the slider

• Report cursor – using the toggle button

When communication with the Graphic server is lost and the system does

not respond to any of your actions on the 3-Plane GRx screen for an

amount of time, an error message posts (Figure 3-3) and the 3-Plane GRx

window closes. You are able to re-enter 3-Plane GRx by clicking the

Graphic Rx icon again, retaining all of your deposited prescriptions.

Modality Worklist -Connect Pro

The Performed Procedure Step (PPS) is a closed loop of information transfer. The loop starts with the HIS/RIS (Hospital Information System and/or the Radiology Information System), when the patient information is transferred to the system after you select the patient from the Schedule window. The PPS is completed when the images are archived and transferred to a PACS.The HIS/RIS is notified that the exam is complete and of the location of images (e.g., at the PACS). ConnectPro Plus is a purchasable option that allows information sharing between the HIS/RIS and the MR system.

Procedure codes can be configured on the system and linked to protocols in the system protocol library. When patient information from the HIS/RIS is called to the system through ConnectPro, the procedure code is part of the information. If the code is linked to a system protocol, that protocol is loaded to the Scan Rx desktop along with the patient information. At the New Patient window in the Scan Rx desktop, two or more numbers entered at the Accession Number or Patient ID text box prompt a search through the ConnectPro Worklist for a match. If a match is found, the patient

information is entered into the Patient Information window.

E3DTOF

FSE_XL

The FSE-XL and FSE-XL-IR with Blood Suppression sequences utilize an inversion recovery preparation pulse to null signal from blood to obtain “black blood” cardiac images.

BloodSupp

The FSE-XL and FSE-XL-IR with Blood Suppression sequences utilize an inversion recovery preparation pulse to null signal from blood to obtain “black blood” cardiac images.

Fastcine

The Wall Motion Sequence is the first series in the cardiac protocol and can

be acquired with a FastCINE SPGR without tags or a FastCINE GRE with

tags. This sequence allows the assessment of the

contractile function of the ventricular wall.

Tagging

The Wall Motion Sequence is the first series of the three sequence cardiac

protocol. It demonstrates functional and anatomical information of the heart.

Tagging can be used with this sequence to better visualize contractile

function.

Sgd perf

DW-EPIZ

iDrive Pro

iDrive™ Pro Plus is an optional software package that provides you with a

flexible environment for using real time imaging. This guide explains how to

image real time with iDrive Pro Plus. It provides key concepts regarding the

acquisition, tools, and image review options available.

Refer to the Real Time Imaging chapter in this guide for additional

information.

iDrive

iDrive™ Pro Plus is an optional software package that provides you with a

flexible environment for using real time imaging. This guide explains how to

image real time with iDrive Pro Plus. It provides key concepts regarding the

acquisition, tools, and image review options available.

Refer to the Real Time Imaging chapter in this guide for additional

information.

SmartPrep 2000 Upgrade

SmartPrep IA (Infusion Angiography) uses a tracking pulse sequence to continuously monitor the MR signal coming from a user prescribed volume of interest in the patient.  As the signal increases, due to an injected contrast material (e.g., gadolinium), the pulse sequence checks to see if the signal amplitude exceeds a contrast dose-dependent threshold.  When the measured signal exceeds the system defined and fixed threshold, the preselected protocol automatically initiated.  To maximize the image contrast at this point, k-space is filled in a centric order.

Probe 2000 Upgrade

(Probe, not 2000 Upgrade)

Map metabolites in the brain from single voxel to 3D chemical shift imaging (CSI). Acquisition, processing and visualization are provided to help characterize brain lesions.

PROBE – PRESS Single-Voxel 
PROBE – PRESS Single-Voxel spectroscopy allows you to non-invasively evaluate the relative concentrations of in-vivo metabolites.  It lets you acquire and display volume localized, water-suppressed H1 spectra in single-voxel mode. This package includes the PROBE-PRESS pulse sequence as well as automated reconstruction, acquisition set-up and graphic prescription of spectroscopic volumes.

PROBE-PRESS and PROBE-STEAM Single-Voxel 
For advanced spectroscopy users, this enables single-voxel capability with both the PROBE-PRESS and PROBE-STEAM pulse sequences.

PROBE 2DCSI 
This capability lets you extend your PROBE-PRESS spectroscopic capabilities to perform 2D CSI acquistions, thereby enabling simultaneous multi-voxel, in-plane acquistions. Post-processing, including the creation of metabolite maps, is automatically generated with the FuncTool Performance Package (included in ScanTools). Signa* HDxt 1.5T supports true, multi-channel PROBE 2DCSI capabilities.

PROBE 3DCSI 
With this capability, you can extend advanced PROBE-PRESS 2DCSI spectroscopic capabilities to include three-dimensional, multi-voxel acquisitions. All post-processing, including the creation of metabolite maps, is automatically generated with the FuncTool Performance Package (included in ScanTools). Signa* HDxt 1.5T supports true, multi-channel PROBE 3DCSI capabilities.

Multi-Nuclear Spectro

The Multi-Nuclear Spectroscopy (MNS) option consists of the software and hardware necessary to acquire,

display, and store 31P signals on the Signa System. MNS uses PROBE 2000 and SAGE 7 software.

PROBE 2000 is an image guided, clinical spectroscopy package which permits the acquisition of volume localized,

water suppressed hydrogen spectra in single- or multi-voxel modes. Automated acquisition setup including the

adjustment of homogeneity and water suppression is a standard part of the package.

PROBE 2000 spectra provide information about the relative concentrations of in vivo metabolites in a noninvasive

manner. This data may be visualized as individual spectra or as spectral images. In addition, the results of

a quantitative analysis of the spectrum may be displayed along with the spectrum using the image browser tools.

SAGE 7 is a spectroscopic processing and display software tool. It provides a powerful toolkit of software

applications to handle a wide variety of tasks associated with spectroscopy data management; including file

handling, display, processing, analysis, storage and hardcopy output.

For more information on MNS, refer to the following manuals:

• MR Excite™ HD Advanced and Multi-Nuclear Spectroscopy operator manual, direction 5123076-100.

• Sage 7 LX Users Guide direction number is 2114789-100.

Functool2

FuncTool is an optional software package. It allows the analysis of MR data sets for post processing. The images can represent changes in image intensity over time, b–value or gradient orientation (diffusion imaging), or frequency (spectroscopy). Results are in the form of graphs and functional maps that can be saved and printed to film or a color paper printer.

Voxtool

Interactive Vascular Imaging

Interactive Vascular Imaging (IVI) is a post processing tool that focuses on the creation of vascular projection images. It creates projection images from a 2D stack or 3D volume of Magnetic Resonance Angiography (MRA) images and allows you to view the data from different angles. IVI uses a ray tracing technique called Maximum Intensity Projection (MIP) to create projection images.

Applications

IVI allows you to modify the volume of interest to remove competing structures, such as fat, to improve the vascular images.

Considerations

• Isometric voxels (cube shaped) and overlapping slices improve the quality of IVI projection images.

• All images must have the same orthogonal scan plane, image center, and pixel size.

• The range of images loaded for IVI post processing must not contain two images at the same location.

• The time stamp on IVI images corresponds to the time of the landmark, not the time the source images

were acquired.

• Close IVI prior to starting Real Time for Fluoro Trigger MRA scans.

Clairview

Spectro Analysis GE

Preformed Procedure Step

High Speed

FGRET Pulse Sequences

Fast Gradient Echo-Echo Train (FGRET) is a Fast Gradient Echo (FGRE)

sequence with an echo-planar (EPI) readout. Features of both the FGRE

and EPI pulse sequences are combined for use in cardiac applications. This

guide explains the pulsing components and timing factors directly related to

the FGRET imaging pulse sequence.

FGRET for Real Time

FGRET for Real Time (FGRET-RT) is a fast gradient echo-echo train

sequence performed using RTCA with Signa iDrive Pro Plus to obtain real

time image data with interactive scan plane manipulation.

FGRET-RT, sometimes called "fluoro" mode, has been developed to

provide real time localization of cardiac anatomy. FGRET-RT is prescribed

in the same manner as any iDrive RTCA series. Once data acquisition is

initiated on the Acquire tab, the data acquisition runs continuously. Your

input is required on the Acquire tab to manipulate the scan plane, but

Acquire tab Image Contrast tools are not available for FGRET-RT. You can

review and save the images in the same fashion as any other iDrive image.

NOTE: Refer to Real Time Imaging chapter in this guide for additional

information on iDrive Pro Plus.

Due to the large ETs used in FGRET-RT, the sequence is susceptible to

ghosting artifacts caused by phase errors. The use of fat saturation (SAT)

and low flip angles may help decrease the blurring that occurs with

FGRET-RT imaging. FGRET-RT also provides auto echo alignment, which

is performed to compensate for these phase errors.

Auto echo alignment is performed each time the scan plane is changed

during the FGRET real time session. As a result, there are some images

acquired as echo alignment is taking place, and these images may exhibit

blurring.

Manual echo alignment, a feature also accessed on the User CVs screen,

can be used to check and adjust the echo alignment prior to starting the real

time data acquisition. Manual echo alignment is not commonly used with

FGRET-RT. Refer to this section (What Do I Need to Know About...) for

details on the User CVs screen, and the Manual Echo Alignment procedure

in the How Do I... section for details on how to perform manual alignment.

There are several factors you should consider when using FGRET-RT:

• As in any iDrive RTCA prescription, only one location is prescribed.

• Spatial SAT pulses are not allowed.

High Resolution Spiral

High Resolution Spiral (Hi-Res Spiral) has been developed to obtain high

resolution images in ultra-fast scan times, in seconds and even

milliseconds. This sequence is well suited for obtaining images of the

coronary arteries.

Hi-Res Spiral acquisitions can be performed in cardiac-gated or non-gated

method. To perform a spiral sequence, either Spiral Gradient Echo (GRE) or

Spiral Spoiled Gradient Echo (SPGR) is selected on the Pulse Sequence

window. Selections on the Imaging Options window further define the type

of data to be acquired.

Real Time Spiral Imaging

Real Time Spiral Imaging is used with the iDrive Pro Plus Real Time

Cardiac Acquisition (RTCA) imaging feature. Real Time Spiral acquisitions

are defined by selecting either Spiral GRE or Spiral SPGR on the Pulse

Sequence window and Real Time on the Imaging Options window. As in

any iDrive RTCA prescription, only one location is prescribed for this

sequence.

The following are key components of Real Time Spiral:

• All coil types are allowed.

• Cardiac Gating/Triggering is not allowed.

– The Real Time Spiral frame-rate is typically described in frames per

second (FPS) providing high-temporal resolution images exhibiting

little or no motion in areas such as the heart.

• Real Time Spiral images are of low spatial resolution, so this sequence

is typically used for localization of anatomy.

• The prescribed number of points affects the maximum number of arms

allowed. Likewise, the number of arms affects the maximum number of

points allowed.

• The Imaging Option SSRF is available with Real Time Spiral.

– SSRF can be used to suppress signal from fat. This can be

especially useful for suppressing the fat that may interfere with

visualization of the coronary arteries.

iDrive Pro Plus

iDrive™ Pro Plus is an optional software package that provides you with a

flexible environment for using real time imaging. This guide explains how to

image real time with iDrive Pro Plus. It provides key concepts regarding the

acquisition, tools, and image review options available.

PROBE 3D Brain

Map metabolites in the brain from single voxel to 3D chemical shift imaging (CSI). Acquisition, processing and visualization are provided to help characterize brain lesions.

Ultrashort TR

T2 Breathhold

Use of breath-hold imaging with a rapid technique helps reduce image noise from respiratory or motion artifacts and enables higher patient throughput and easier examination of uncooperative, medically unstable, or claustrophobic patients.

The breath-hold fast-recovery technique makes use of additional radio-frequency pulses after the final acquisition window to drive the recovery of longitudinal magnetization.

SSFSE MRCP

Single Shot Fast Spin Echo (SSFSE) and Single Shot Fast Spin Echo-Inversion Recovery (SSFSE-IR) use an RF pulse design that allows for significantly shorter echo spaces (ESP). The minimum TE no longer equates to the ESP.

The scan time for an SSFSE acquisition = (Acq rate + SAR delay) x (# of slices).

• The Acquisition Rate is the product of the ETL times the echo spacing.

T1 Breathhold

Line Scan Diffusion Imag

A novel line scan diffusion imaging sequence (LSDI) is intro-

duced. LSDl is inherently insensitive to motion artifacts and

high quality diffusion maps of the brain can be obtained rap-

idly without the use of head restraints or cardiac gating. Re-

sults from a stroke study and abdominal diffusion images are

presented. The results indicate that it is feasible to use the

LSDl  technique  for  clinical evaluation  of  acute  ischemic

stroke.  In  contrast  to echo-planar diffusion imaging,  LSDl

does not require modified gradient hardware and can be im-

plemented on conventional scanners. Thus, LSDl should dra-

matically increase the general availability of  robust clinical

diffusion imaging.

Image Filter

ACGD Plus

Probe 3D Prostate

Fluoro- triggered MRA

Fluoro Trigger is an option available for Fluoro Triggered Magnetic

Resonance Angiography (FT MRA) imaging. This option must be used in

order to perform a FT MRA sequence.

MRCP3

An axial or coronal MR Cholangiogram (MRCP) scan can be acquired with

a 3D FIESTA pulse sequence. This sequence provides enhanced contrast

and temporal resolution for imaging of the bile ducts. Figure 5-5 shows an

image of the bile ducts acquired with a 3D FIESTA sequence. Table 5-7

gives you a brief look at an example protocol. The parameters may vary

depending on your needs.

PSIR

Dynamic R1

Dynamic R1 can be used with a 3D FSE pulse sequence to

increase the SNR for acquiring large FOV images. If during

manual prescan, the R1 of the imaging volume is less than 6,

consider using Dynamic R1.

Fiesta 2D

The Fast Imaging Employing STeady-state Acquisition (FIESTA) is a fully

balanced steady-state coherent imaging pulse sequence designed to produce

high signal-to-noise ratio (SNR) images at very short sequence times (TRs).

This chapter explains the pulsing components and timing factors directly related

to the 2D FIESTA imaging pulse sequence. It also explains the concepts and the

step-by-step instructions to help you learn how to:

• Optimize 2D FIESTA Images

• Prescribe a 2D FIESTA Sequence

In addition, this chapter answers the following questions:

1. What is FIESTA?

2. When would I use a 2D FIESTA pulse sequence?

3. What are the imaging characteristics of a 2D FIESTA pulse sequence?

4. Which imaging options can be used with a 2D FIESTA pulse sequence?

5. What are some applications for the 2D FIESTA sequence?

Fiesta 3D

The Fast Imaging Employing STeady-state Acquisition (FIESTA) sequence

is a fully balanced steady-state coherent imaging pulse sequence designed

to produce high SNR images at very short TRs. The pulse sequence uses

fully balanced gradients to re-phase the transverse magnetization at the

end of each TR interval. For very short TR sequences, the resulting signal

intensity is independent of TR and related to (T2/T1). More importantly, the

short TRs are essential to maintain spin phase coherence. Phase

coherence is required to maintain coherent transverse magnetization and

eliminate artifacts generated by magnetic susceptibility induced phase

shifts. Thus, the advantages of FIESTA can only be realized with very short

TR where (TR << T2) and (TR << 1/b) where b is the local frequency shift

caused by inhomogeneity.

NOTE: << indicates “much less than”.

Tissue contrast is generated based on the ratio of the spin-spin relaxation

time (T2) and the spin-lattice relaxation time (T1). Hence, the pulse

sequence accentuates the contrast of spins with high T2/T1 ratios (such as

cerebral-spinal fluid, water, and fat) while suppressing signal from tissues

with low T2/T1 ratios (such as muscle and myocardium).

The 3D FIESTA sequence can be used for whole body imaging and can be

used in clinical applications that benefit from the differentiation of contrast

between tissues of low T2/T1 ratios (low signal intensity) and high T2/T1

ratios (high signal intensity). The 3D FIESTA technique can be useful for,

but not limited to, imaging structures in motion such as abdominal imaging

of the bile ducts or for rapid acquisition of static structures with high spatial

resolution such as the cochlea or joint imaging.

Delayed Enhancement

On delayed enhancement MR images, scarring or fibrosis appears as an area of high signal intensity that is typically subendocardial or transmural in a coronary artery distribution.

The technique for delayed enhancement MR imaging involves intravenous infusion of gadolinium chelate contrast material (0.1–0.2 mmol/kg) followed 10–30 minutes later by a cardiac-gated T1-weighted pulse sequence. Imaging too early (eg, less than 5 minutes after the initial contrast material injection) results in reduced contrast difference between infarcted and normal myocardium because insufficient contrast material has been washed out of the normal myocardium. This can lead to overestimation of the infarcted region (2). However, imaging too late (eg, more than 30 minutes after initial contrast material injection) may result in excessive washout of the contrast agent from the infarcted tissue and poor signal-to-noise ratio.

The typical pulse sequence for myocardial delayed enhancement is a segmented inversion-recovery–prepared fast gradient-echo sequence. Imaging occurs over nine to 12 heartbeats, in a breath hold. In patients who have difficulty holding their breath, navigator-assisted free-breathing techniques can be used without breath holding.

Asset

Array Spatial Sensitivity Encoding Technique (ASSET) is an option that

enables faster scanning when using the Torso Phased Array (PA) coil. It is

designed to achieve shorter breath-held acquisitions, thus making it easier

for your patients experiencing difficulty holding their breath for long periods

of time.

ProbePRESS

# PROBE - Allows up to an Entire Brain Evaluation

From single pixel thru 3D whole brain test, MR Spectroscopy allows multi-voxel evaluation of the entire brain to help characterize lesions. PROBE 3D CSI Software is an addition to the PROBE 2000 package that enables 3D Multi-Voxel Proton Spectroscopy.

3D Multi-Voxel capabilities provide a significant enhancement to a standard Proton Spectroscopy by enabling improved anatomic coverage, SNR and resolution. PROBE is fast and easy to use, producing hi-resolution spectra in as little as 10 minutes.

Parametric metabolite maps can be created and then overlaid over the anatomical MR images for improved diagnostic confidence. Spectroscopy image processing can be performed on the operators console or on an Advantage Windows workstation.

CREATIV

Navigator

Navigator is a cardiac option that enables you to perform navigated,

free-breathing, coronary artery acquisitions. The Navigator Imaging Option

is used with a 3D enhanced Fast Gradient Echo (GRE) pulse sequence to

acquire the images.

Diffusion Tensor

Diffusion Tensor Imaging (DTI) is the newest addition to the Echo Planar

pulse sequence family. It is used to visualize the white matter fiber tracts in

the central nervous system (CNS) and provides information on molecular

diffusion of biological tissues.

3DFRFSE

3D Fast Recovery Fast Spin Echo-Accelerated (FRFSE-XL) pulse sequence is built on the foundation of the 3D FSE pulse sequence. It has been developed primarily to enable high-resolution images for MR cholangiopancreatography (MRCP) studies.  The fast recovery feature is designed to enhance the intensity of fluids that have long T2 relaxation times, while using a shortened TR time. The shortened TR time makes the high resolution images practical.  A sequential view ordering scheme can be used with the 3D FRFSE-XL sequence. This acquisition mode is called the Single Shot mode. It fills an entire ky (phase) and kx (frequency) plane of k-space in a single shot for each value of kz (slice) and is not interleaved with kz.

Asset Plus

(this isn't Plus)

Array Spatial Sensitivity Encoding Technique (ASSET) is an option that

enables faster scanning when using the Torso Phased Array (PA) coil. It is

designed to achieve shorter breath-held acquisitions, thus making it easier

for your patients experiencing difficulty holding their breath for long periods

of time.

Fat Water 2001

Rapid Frame

TRICKS

Elliptic Centric-TRICKS (Time Resolved Imaging of Contrast Kinetics) produces the following:

• CE MRA high spatial and temporal resolution images.

• A mask acquisition used to produce automatically subtracted source images.

• Collapsed images from each temporal output phase.

TRICKS high temporal resolution is achieved by dividing the 3D k-space into a number of segments from the center of k-space out (A to D in illustration below). Views are acquired in elliptic centric order and the rate of sampling is varied such that the center of k-space is sampled more often than the outer regions. When the center of k-space is sampled more frequently than other regions then the time period from one phase to the next is shortened. The end result is that the contrast kinetics/flow is subdivided into more phases with TRICKS than with other PSDs and therefore the temporal resolution is shorter than other PSDs.

FIESTA-c

The Fast Imaging Employing STeady-state Acquisition (FIESTA) - C (Cycled phases) sequence is similar to 3D FIESTA in that it is a fully balanced steady-state coherent imaging pulse sequence designed to produce high SNR images at very short TRs. FIESTA-C acquires multiple images with a cycling of the RF phase to reduce banding artifacts and improve image homogeneity. FIESTA-C collects two images at different receiver phases and then generates a MIP image across the phase cycles for each location in the 3D acquisition. It is these MIP images that are saved in the Browser. The RF phase cycling allows for use of longer TR values (> than 4-5 ms) while maintaining the same

benefits of FIESTA excellent contrast differentiation between structures of low signal intensity and high signal intensity.

BREAST 2

Propeller DWI

(3.0)

Combat patient motion, flow and susceptibility artifacts anywhere.

PROPELLER 3.0 is designed to reduce the effect of patient voluntary and physiologic motion (breathing, flow, peristalsis), and reduce magnetic susceptibility artifacts. It helps generate consistently good, diagnostic quality images even for challenging patients and difficult-to-image anatomies.

PROPELLER 3.0 uses a radial k-space filling pattern that, compared to the Cartesian method, is inherently less sensitive to motion such as CSF and blood flow, breathing, patient tremor, or voluntary movements. In addition, a sophisticated motion correction, post-processing algorithm is deployed to further reduce effects of rigid motion. The oversampling of the k-space center typical for radial k-space filling yields increased signal-to-noise ratio (SNR) and high tissue contrast. Scan time is the same as conventional techniques, and PROPELLER allows free-breathing acquisition.

PROPELLER 3.0 is enabled for T1 FLAIR, T2, T2 FLAIR imaging in all planes, axial diffusion weighted imaging for brain, T2-weighted imaging for cervical spine, Body, and T2/PD-weighted imaging for MSK.

3D Delayed Enhancements

3D Fat Sat FIESTA

3D Fat Sat FIESTA technique for free-breathing coronary artery imaging. It provides exquisite anatomical detail and the volumetric acquisition makes it easier to assess structures that have a tortuous course.

MRAppStudio

BrainWave RT

(I put same for PA and HW)

BrainWave™ is the brain mapping software on the AW system used to record the data acquired during an fMRI exam. It integrates the sensor device with the software. BrainWaveSO is the same brain mapping software with the same applications as BrainWave except it does not include

integrated sensor device information.  How can you tell if your system is operating BrainWave versus

BrainwaveSO software? The button on the AW Browser is labeled either BrainWave or BrainWaveSO. In addition, the menu titles display either BrainWave or BrainWaveSO.  This software is used to start your functional brain acquisition and visualize brain maps. The BrainWave software is divided into four main windows, which can be accessed from the Main Menu (Figure 8-5).

• Acquire fMRI Scans

• Visualize Maps

• Re-Process fMRI Scans

• Manage System

Each window is equipped with several functions, as well as the BrainWave or BrainWaveSO Message and Warning text areas. The BrainWave Message area displays information regarding the system status. For example, the system may display messages such as, “BrainWave is in idle mode” or “BrainWave is now importing data to the AW Browser.” The top of the message area always displays the software version. The Warnings area displays information regarding when the system is not responding normally.  For example, the system may display a message such as, “BrainWave is unable to import data to AW Browser.” After a warning has been posted, you should check to see if the problem can easily be resolved or contact your service engineer to correct the problem.

BrainWave PA

BrainWave™ is the brain mapping software on the AW system used to record the data acquired during an fMRI exam. It integrates the sensor device with the software. BrainWaveSO is the same brain mapping software with the same applications as BrainWave except it does not include

integrated sensor device information.  How can you tell if your system is operating BrainWave versus

BrainwaveSO software? The button on the AW Browser is labeled either BrainWave or BrainWaveSO. In addition, the menu titles display either BrainWave or BrainWaveSO.  This software is used to start your functional brain acquisition and visualize brain maps. The BrainWave software is divided into four main windows, which can be accessed from the Main Menu (Figure 8-5).

• Acquire fMRI Scans

• Visualize Maps

• Re-Process fMRI Scans

• Manage System

Each window is equipped with several functions, as well as the BrainWave or BrainWaveSO Message and Warning text areas. The BrainWave Message area displays information regarding the system status. For example, the system may display messages such as, “BrainWave is in idle mode” or “BrainWave is now importing data to the AW Browser.” The top of the message area always displays the software version. The Warnings area displays information regarding when the system is not responding normally.  For example, the system may display a message such as, “BrainWave is unable to import data to AW Browser.” After a warning has been posted, you should check to see if the problem can easily be resolved or contact your service engineer to correct the problem.

BrainWave HW

BrainWave™ is the brain mapping software on the AW system used to record the data acquired during an fMRI exam. It integrates the sensor device with the software. BrainWaveSO is the same brain mapping software with the same applications as BrainWave except it does not include

integrated sensor device information.  How can you tell if your system is operating BrainWave versus

BrainwaveSO software? The button on the AW Browser is labeled either BrainWave or BrainWaveSO. In addition, the menu titles display either BrainWave or BrainWaveSO.  This software is used to start your functional brain acquisition and visualize brain maps. The BrainWave software is divided into four main windows, which can be accessed from the Main Menu (Figure 8-5).

• Acquire fMRI Scans

• Visualize Maps

• Re-Process fMRI Scans

• Manage System

Each window is equipped with several functions, as well as the BrainWave or BrainWaveSO Message and Warning text areas. The BrainWave Message area displays information regarding the system status. For example, the system may display messages such as, “BrainWave is in idle mode” or “BrainWave is now importing data to the AW Browser.” The top of the message area always displays the software version. The Warnings area displays information regarding when the system is not responding normally.  For example, the system may display a message such as, “BrainWave is unable to import data to AW Browser.” After a warning has been posted, you should check to see if the problem can easily be resolved or contact your service engineer to correct the problem.

Propeller T2

PROPELLER HD yields high-quality brain imaging that is extremely resistant to motion artifacts, especially useful in cases where patients are difficult to image due to movement, including Parkinson’s patients who suffer from uncontrollable patient motion and children who do not respond to sedation.  PROPELLER™ HD imaging is an advanced technology that combines the advantages of fast spin echo (FSE) and radial data acquisition to create high-quality  T2-weighted and diffusion-weighted brain images. PROPELLER HD improves contrast-to-noise ratio by 20 to 30 percent,  significantly reduces the tissue-to-air and tissue-to-metal image distortions of diffusion-weighted imaging and dramatically

reduces the sensitivity of image quality to patient motion.  UPMC uses both PROPELLER HD and FSE imaging techniques for T2-weighted imaging on essentially all brain MR exams performed on the Signa 1.5T MR imaging system.

3T Restricted Apps

2D Fat Sat FIESTA

(same as 2D Fiesta)

The Fast Imaging Employing STeady-state Acquisition (FIESTA) is a fully

balanced steady-state coherent imaging pulse sequence designed to

produce high signal-to-noise (SNR) images at very short sequence times

(TR). This guide provides you with information on the pulsing components

and timing factors directly related to the 2D FIESTA imaging pulse

sequence.

MR-Echo

MR-Echo provides the ability to acquire real time images of the heart using a FIESTA (bright blood) pulse sequence.

There are two MR-Echo modes of operation accessed from the two tabs on the Scan desktop:

• RealTime Parameters tab provides a high frame rate acquisition mode for localization and qualitative

ventricular function assessment. This mode works with and without ECG gating. The images can be saved

to the database (Browser).

• Scan and Save Parameters tab provides a high frame rate acquisition mode for high image quality breath

hold cardiac images that are added to the exam database. This mode only functions with ECG gating. MREcho

images can be exported into cardiac analysis tools such as ReportCARD.

Temporal resolution is referred to through out this section. Temporal resolution refers to the time it takes to

acquire each phase of the cardiac cycle. The lower the temporal resolution number, the better the temporal

resolution. Typically, a temporal resolution of 80 ms or less provides sufficient image quality for qualitative or

quantitative myocardial function analysis.

8 Channel System

16 Channel System

Lava

LAVA Flex generates four image contrasts: Water, Fat, In-Phase and Out-of-Phase in a single breath-hold.

LAVA Flex is a 3D, FSPGR imaging technique that generates water-only, fat-only, in-phase and out-of-phase echoes in one single acquisition that is typically completed in one 20 second breath-hold. This technique provides excellent homogeneous fat suppression over the entire field of view, including areas that are difficult to image using conventional fat suppression, due to magnetic susceptibility effect.

All four contrasts are perfectly coregistered because they are acquired in the same scan. The water-only contrast differs from a conventional fat suppressed image in that an inversion prep pulse is not applied for fat suppression. In fact, the fat information is removed, leaving a water-only image that may potentially be used in place of a LAVA type image. LAVA Flex uses ARC*, a 2D self calibrated, parallel-imaging technique that allows for acceleration in both phase and slice directions for supported coils.

FiberTrak

FiberTrak

FiberTrak is an advanced post-processing tool designed to generate eigenvector information from the diffusion tensor acquisition and processing. Using a robust and efficient seeding process, three dimensional color renderings of the diffusion along white matter tracts can be then generated and coregistered with anatomical images such as BRAVO.  FiberTrak maps can be stored in DICOM format. You can also coregister tractography with fMRI activation maps overlaid on an anatomic 3D image using BrainWave Fusion.

FLAIR3D

The T1 FLAIR pulse sequence is designed to reduce the T1 weighted imaging scan time, as compared to T1 weighted Spin Echo (SE) imaging scan times. It also is designed to achieve better tissue contrast-to-noise (CNR) and tissue SNR.  You can activate the T1 FLAIR sequence by selecting T1 FLAIR in the Pulse Sequence window. This Pulse Sequence Database (PSD) now operates in both the Picture This and Site Protocol modes. T1 FLAIR images are annotated as “t1flair/90” in the lower left corner of images. NOTE: The Site Protocol mode allows you to adjust prescription parameters and save the protocol in your site library.  Pausing a FLAIR acquisition during scanning may cause image artifacts and should be avoided.

Multi-Phase (Variable delays)

The Multi Phase FGR-ET Sequence is the second series in the cardiac protocol. This multi-slice sequence is typically repeated to acquire multiple phases of the cardiac cycle.

Dual Array

Timeout Application

Blood Flow and Volume Measurements

Time Course

# FGRE Time Course

Excellent coverage and multiplane visualization designed to help you perform time course studies with clinical confidence.

The first choice for stress studies, FGRE Time Course delivers excellent temporal and spatial resolution images in multiple planes, to help capture difficult to see pathologies, like the apical thrombus (above)..

Benefits include:

- Multiplanar acquisition for simultaneous short and long axis visualization

- Excellent temporal resolution (3-4 slices per RR)

- Resistant to off-resonance & eddy current effects

- Excellent for stress study on stunned vs. infarcted myocardium

- FGRE Time Course deploys single shot FGRE with interleaved slice acquisition. The background suppression options include notch-selective inversion pulse for high temporal resolution and saturation pulse resulting in uniform fat suppresion in multiplanar acquisition.

- Shorter excitation RF pulse (200 μs) accelerated with ASSET.

SWIFT

Sweep Imaging with Fourier Transformation..

T2MAP

Cartigram (T2 Map) is an optional software package used to noninvasively detect changes in the collagen component of the extracellular matrix of cartilage. Cartigram acquires multiple scans at each slice location; each set of scans has a unique TE value. The acquired data can then be processed in FuncTool to produce color maps that demonstrate subtle changes in the cartilage structure that are not visible of gray scale MR images. Collagen in the radial zone of articular cartilage close to the subchondral bone is more highly ordered with small quantities of mobile water. By contrast, collagen in the transitional zone is more randomly oriented with increased water mobility and longer T2 values. Studies have shown that an increase in water mobility (increase in T2 relaxation time) may be associated with damage to the collagen matrix.

2D MERGE

High-resolution spine and MSK images with excellent tissue contrast.

Multiple Echo Recombined Gradient Echo (MERGE) uses multiple echoes to generate high-resolution images of the C-spine with excellent gray-white matter differentiation. By combining early echoes with high SNR and late echoes with improved contrast, the result is improved cord contrast within the spinal column.

3D MERGE also offers excellent SNR and fat-saturation capabilities to provide high resolution, isotropic T2* weighted images of the extremities (hand, wrist, knee, ankle, and shoulder) with optimized contrast between the ligaments and soft tissue.

- Fast, high resolution axial volumetric scan

- Selective water excitation for fat-free imaging1

- Intuitive reformat in virtually any plane

LAVA-XV

SMART-P

BRAVO

Excellent T1w contrast between grey and white matter. Fast, high resolution 3D acquisition.

Brain Volume imaging is a high-resolution 3D imaging technique designed to produce heavily T1-weighted isotropic images of the brain.

BRAVO deploys 3D IR-prepared FSPGR acquisition to produce isotropic T1-weighted volumes.  The center of k-space is oversampled and serves as the calibration data for the parallel-imaging reconstruction. Innovative 2D ARC parallel imaging helps reduce scan time and minimize parallel imaging artifacts.

- Scan once, reformat into any plane

- Reliably visualize even small and subtle lesions

- Automated, rapid, whole brain volume exam

- Excellent T1w optimized for grey-white matter contrast also for BOLD studies

Breast Single Voxel Spectroscopy

Single Voxel Spectroscopy optimized for noninvasive breast tissue assessment.

A TE-averaged PRESS (Point RESolved Spectroscopy) acquisition that provides the necessary biochemical information to help characterize breast tissue.

- Literature shows that spectroscopy has a useful role in breast imaging.

- Assessing presence of choline can help increase specificity in lesion assessment.

- An easy to use1, breast-specific single-voxel MRS tool for additional lesion characterization.

- BREASE can also help assess early treatment response.

Image shown above courtesy of HIA Laveran

1. Compared with conventional MRS

COSMIC

Fast, high resolution axial volumetric scan.

This is a 3D sequence used to image the axial c-spine. COSMIC (Coherent Oscillatory State Acquisition for the Manipulation of Imaging Contrast) uses a modified fast GRE pulse sequence with steady-state free precession segmented multi-shot centric k-space acquisition. This improves the CNR and SNR of c-spine structures including the spinal cord, vertebral disks, nerve root canal and contrast between CSF and nerve roots.

- Fast, high resolution axial volumetric scan with ability to reformat to different planes

- High SNR

- Excellent contrast among disc, muscle and bony structures, CSF, nerve tissue

32 Channel System

32-channel torso coil from GE Healthcare geometrically overlaps coil elements to achieve signifcant SNR and parallel imaging improvements over 8 channel technology. At high field MR such as 3.0T, there is an abundance of SNR that may be utilized to trade off scan time; therefore, new coil

technologies that maximize parallel imaging capabilities are imperative.

Integrated Patient Comfort Module

Brainwave Fusion

BrainWave Fusion is an optional package that provides the ability to fuse high-resolution anatomical images with fMRI activation maps and diffusion tensor tractography maps. This package is useful for evaluating the spatial relationship between activation patterns, fiber tracts, and underlying anatomy and pathology.

Hdx Key

IDEAL

IDEAL provides robust, homogenous fat/water separation – even in difficult-to-scan anatomies.

## IDEAL

In MR imaging, magnetic susceptibility effect can obscure or mimic pathology – especially in challenging anatomies. IDEAL provides robust, homogenous fat/water separation even in difficult to scan anatomies. It generates four different contrasts – water-only, fat-only, in-phase, and out-of-phase – from one single acquisition. Robust, homogenous fat suppression and additional tissue insights aid in more-confident diagnoses and eliminate the need for a separate fat saturation scan.

Neuro MRI
IDEAL reduces chemical shift artifacts in images of difficult to scan anatomies, such as spine, orbits, skull base, and neck.

MSK MRI
IDEAL helps preserve image integrity and suppress fat also in areas of high magnetic susceptibility.

Body MRI
IDEAL delivers robust, consistent fat/water separation, and additional tissue information, such as presence and distribution of fat, to aid more confident diagnosis.

3D Dual Echo

# 3D Dual Echo - Uncompromised image quality in a single breath-hold

3D Dual Echo is a new sequence that produces in-phase and out-of-phase images in the same acquisition within a single breath-hold. It ensures perfect registration between corresponding images, and enhances small lesion detection and visualization.

# 3

ARC

Cube T2

Scan once, then reformat to any plane with high sub-millimeter resolution.

Cube* 2.0 replaces several 2D slice acquisitions with one single 3D volume scan. Sub-millimeter isotropic Cube volume data can be easily reformatted into any plane—without gaps, and with the same resolution as the native plane. The new self-calibrating, parallel-imaging engine ARC helps to speed up the sequence while minimizing artifacts. And you'll appreciate the high tissue contrast.

- Neuro MRI. Cube 2.0 acquires contiguous, sub-millimeter isotropic 3D data that can be easily reformatted into any plane, thus replacing multiple 2D scans. SNR-rich ultra-thin slices help you visualize even small and subtle lesions without partial-volume averaging effect. Cube is enabled for T1, T2, T2 FLAIR, and PD contrasts.

- MSK MRI. Cube 2.0 generates contiguous, sub-millimeter isotropic 3D data that can be later easily reformatted into any plane. You can now easily examine some tissue structures that are best seen in oblique views. SNR-rich, ultra-thin slices help eliminate partial-volume averaging effect, and visualize even small and subtle lesions.

- Body MRI. Cube 2.0 acquires contiguous, sub-millimeter isotropic 3D data that can be easily reformatted into any plane, thus replacing multiple 2D scans. SNR-rich, ultra-thin slices help you visualize even small and subtle lesions without partial-volume averaging effect.

T2Star Weighted Angioraphy

T2 star weighted angiography (SWAN) is a new method for susceptibility-weighted imaging with short acquisition times.

Multi-Echo fgre

Emulation Override

Inhance 2D inflow

Inhance 3D Velocity

Inhance 3D inflow IR