## Project Info
Project Type | Location | Project ID | Customer | Site Name | Site Address | Shipping Address | Contract Type
Install | Fixed Site
De-install | Mobile Site
Upgrade | Relocatible
Special project | Warehouse
Vendors | Contact Name | Phone number | Notes
Site Manager
Contractor
Project Manager
Shipper
Rigger
Other
Modality | Size | Type | Software | Options | Coils | Equipment
MRI | 0.35T | S1 | 1x | Echo Planar Imaging | CRM
CT | .5T | S2 | 2x | Fast Gradient Echo | BRM
PET | .7T | S3 | 3x | CINE | TRM
PET/CT | 1.0T | S4 | 4x | Fast Spin Echo & FLAIR | XRM
X-Ray | 1.5T | S5 | 5x | Time of Flight | Breast Coil
Mamo | 3T | SX | 6x | Phase Contrast Vascular | Open Breast Coil
Ultra Sound | 7T | CX | 7x | Research | Vibrant Bilateral Breast Array Coil
Dexa | H1 | LCC | 8x | Proactive Service | Cardiac Phased Array Coil
| H2 | TWIN | 9x | Research PSD | Channel Cardiac Array Coil
| H3 | MR40 | 10x | Spectroscopy/ PROBE | Extremity Coil
|  | MR750 | 11x | Videoconferencing | Signa Coil
|  | LXI | 12x | SGD_EchoSpeed | Knee Phased Array Coil
|  | DXI | 14x | DW EPI | Lower Leg Array Coil
|  | FXI | 15x | FLAIR EPI | HD Knee Array Coil
|  | LS | 16x | SPECIAL | T/R Knee/Foot Coil
|  | LS Ultra | 20x | Smart Prep | High Resolution Wrist Coil
|  | LS Plus | 21x | SSFSE | General Pupose Flex Coil
|  | LS Xtream | 22x | Three Plane Localizer | Channel Flex Phased Array Coil for Signa Excite HD
|  | VCT |  | Modality Worklist | MNS Phosphorus Flex Coil
|  | NXI |  | E3DTOF | Head Coil Tuning Check (Classic)
|  | DLS |  | FSE_XL | Signa Infinity Split Head Coil
|  | DST |  | Bloodsupp | Channel High Resolution Brain Array Coil
|  |  |  | Fastcine | Quadrature NV Coil
|  |  |  | Tagging | Quadrature NV Array Coil
|  |  |  | sgdperf | HD NV Array Coil
|  |  |  | DW-epiz | Phased Array NV Coil
|  |  |  | iDRIVE Pro | Open NV Coil
|  |  |  | iDrive | Channel NV Array Coil
|  |  |  | Smart Prep 2000 Upgrade | Flow 7000 Phased Array Peripheral Vascular Coil
|  |  |  | Probe 2000 Upgrade | CV/i Peripheral Vascular Coil
|  |  |  | Multi -Nuclear Spectro | Phased Array Shoulder Coil
|  |  |  | Functool2 | Shoulder Coil
|  |  |  | Voxtool | MRID Phased Array Shoulder Coil
|  |  |  | Interactive Vascular Imag | Mark 9000 Phased Array Shoulder Coil
|  |  |  | Clairview | Quadrature Cervical Spine Coil
|  |  |  | Spectro Analysis GE | Anterior Neck Coil
|  |  |  | Preformed Procedure Step | Posterior Neck Coil
|  |  |  | High Speed | Quadrature T/L Spine Coil
|  |  |  | fgret | CTL Phased Array Coil
|  |  |  | fgret Real Time | Profile 7000 Phased Array Volume Neck Coil
|  |  |  | Spiral Hi-Res | CTL Phased Array Coil
|  |  |  | Spiral Real Time | Premier 9000 8-Channel CTL Array Coil
|  |  |  | iDrive Pro Plus | 8-Channel CTL Spine Coil
|  |  |  | PROBE 3d Brain | Signa Horizon Pelvic Phased Array Coil
|  |  |  | Ultrashort TR | Torso Phased Array Coil
|  |  |  | T2 Breathhold | Prima 9000 8-Channel Body Array
|  |  |  | SSFSE MRCP | ATD Torso (ATD-T) Assembly
T1 Breathhold
Line Scan Diffusion Imag
Image Filter
ACGD Plus
Probe 3D Prostate
Fluoro-triggered MRA
MRCP3
PSIR
Dynamic R1
Fiesta 2D
Fiesta 3D
Delayed Enhancement
Asset
ProbePRESS
CREATIV
Navigator
Diffusion Tensor
3DFRFSE
Asset plus
Fat Water 2001
Rapid Frame
TRICKS
FIESTA-c
Breast 2
Propeller DWI
3D Delayed Enhancements
3D Fat Sat FIESTA
MR AppStudio
Brain Wave RT
Brian Wave PA
Brain Wave HW
Propeller T2
3T Restricted Apps
2D Fat Sat FIESTA
MR-Echo
8 Channel System
16 Channel System
Lava
Fiber Tank
FLAIR3D
Multi-Phase (variable delays)
Dual Array
Timeout Application
Blood Flow and Volume Measurements
Time Course
SWIFT
T2MAP
2D MERGE
LAVA-XV
SMART-P
BRAVO
Breast Single Voxel Spectroscopy
COSMIC
32 Channel System
Integrated Patient Comfort module
Brainwave Fusion
Data Sheets | Tools/Kits | Moving equipment | Supplies | Parts | Pads/Straps | Phantoms | Extras
System inspection Report | Kit001 Magnetometer | Large forklift | Pallets | See FRUs |  | DQA | Main Disconnect
Install Check List | Kit002 de-ice kit | Small forklift | Large boxes |  |  | ACR | DC Lighting
De-install check List | Kit003 Lakeshore meter kit | Pallet jack | Medium boxes |  |  | Head TLT | Variable lighting resistor
Iso/grad cal data sheet | Kit004 Lakeshore kit | Orange bars | Small boxes |  |  | Body Loader | Chiller
Ramp Up data sheet | Kit005 Ramp probe kit | Table wheels | Large bubble wrap |  |  | LVshim | AC unit
Ramp down data sheet | Kit006 short fill line helium kit | Hydraulic cart | Small bubble wrap |  |  | Quick shim | Race ways
Supercon Shim data sheet | Kit007 ramp and shim cable kit | Scissor lift | Shrink wrap |  |  | 100mm ball | Raised flooring
Grafidy data sheet | Kit008 grafidy kit 1.5T | Boom truck | Tape gun |  |  | CTL Phantoms | Transformer
SNR data sheet | Kit009 cold head change kit | Large craine | Tape |  |  | Bottle Phantom | UPS
PM data sheet | Kit010 CT hoist kit | Small craine | Nitrogen liquid |  |  |  | RDU
Mobile check list | Kit011 MR ramp down kit | Tow truck | Helium gas |  |  |  | Magmon
Fixed site check list | Kit012 shim cable kit | Machine Skates | Helium liquid |  |  |  | Camera
| Kit013 MR ramp supply | CT gantry dollies |  |  |  |  | Injector
| Kit014 MR Ramp and Shim supply | PET gantry dollies |  |  |  |  | Pulse ox
| Kit015 MR Shim supply | Table dollies
Kit016 TPSRF service kit
Kit017 Grad calibration phantom kit
| Kit018 non magnetic gradient crimpper |  |  |  |  | Water phantom
| Kit019 dynamic disable switch tool |  |  |  |  | Small Polly phantom
| Kit020 coil tuning extension cable |  |  |  |  | Medium Polly phantom
| Kit021 Max RF Dummy load sample |  |  |  |  | Large Polly phantom
Kit022 Body coil rotator kit
Kit023 3M service vacuum
Kit024 high pressure regulator
Kit025 1.5T measurement tool
Kit026 manual lake shore measurment
Kit027 ramp prob kit with S series and LCC
Kit028 pressure vacuum kit
Kit029 1.5T RF dummy load
Kit 030 MRI hero kit
Kit031 RDU netbook
Kit032 body coil change kit
Kit033 Hioki meter kit
Kit034 shim camera alignment
Kit 035 Shim camera alignment mod
Kit036 plotting fixture kit
Kit037 De-Installation kit
Kit038 Magnet skate jack kit
Kit039 Nitrogen fill kit
Kit040 HFO ramp cable kit
Kit041 LCC magnet recon temp meter
Kit042 .7 T passive shim disk kit
Kit043 RF door seal repair kie
Kit044 HFO passive shim fixture kit
Kit045 Gauss Meter kit
Kit046 crimp kit
Kit047 .7T graffidy kit

## De-install
De-Install Template
Project ID | Job | MRI saving magnet Tasks | Estimated hrs | Actual hrs | Initals | CT Tasks | PET Tasks | Estimated hrs | Actual hrs | Initals
Project Name
Pre-site Inspection
|  | Complete system inspection report |  |  |  | Schedule inspection
|  | Verify helium level and pressure (schedule fill if needed) |  |  |  | Verify equipment and Gantry exit paths
|  | Verify magnet and equipment exit paths and sizes |  |  |  | Verify system operation, tube hrs, options and any errors
|  | Schedule start date and rig out date of system |  |  |  | Verify system against sales quote
Schedule rig out date
Verify all dollies and moving equipment shipped
Verify packing supplies and or U-line Shipment
Verify fork Lift and or tow truck for Gantry and equipment
Verify truck scheduled and rig out time
| Site readiness | Verify fork lift and or craine for magnet |  |  |  | Verify you have all tools and equipment needed
Verify fork lift and pallet jack for pallets
Verify access and work hours to get it done
Identify any posible issues and work through them
Verify system operation, tube hrs, options and any errors
Verify save state, options MODs or DVDs and manuals
Lock out tag out procedure with facility
Disconnect all cables from console area
Disconnect all cables from scan room
Disconnect all cables from equipment room (if applicable)
| Pre Shipping |  |  |  |  | Warp and package all cables and equipment
|  | Verify U-line and packing shipment |  |  |  | Separate any extras and optional equipment
Verify ramp down kit, bars, helium and or tools arival
|  | Verify magnet and system truck scheduled |  |  |  | Cordinate equipment removal and gantry rig out
|  | Verify scrap truck (if scrapping magnet) |  |  |  | Verify gantry and table are locked and lowered onto the floor
|  | Verify body coil pallet or tools to buidl pallet |  |  |  | Verify all pallets and equipment is secured in shipping container
Verify magent and equipment delivery Location
Sign off on rigging equipment and shipper
Verify delivery location and arival date
Verify complete system was received and in good condition
System Functional Checks
CTL Phased Array Scan - Setup, Landmark, then use Auto Prescan
Verify Options Against Site Evaluation- Discrepancies, Call Don or Russ
Verify Coil List Against Site Evaluation – Discrepancies Call Don or Russ
Review Error Log - Go back two weeks, Gradient RF or T/R driver faults, Call Don or Russ
Run saveINFO, save gui - Under Guided Install
Make sure to remove all CD's, MOD's, and DAT's
Run Dog House all the way to the Back of the Magnet
Shut down the system, Turn off the breaker at the A1 for the MRI PDU
Jumper out the E-stop loop in the A1 (Run a wire from terminal 3 to 7)
Make Sure That Coldhead Stays Running; Unless You are Scrapping the Magnet
De-commission system
Check Vent Pipe integrity and exit path, warn everyone in area, Check Helium and Pressure
Setup Ramp down Kit and ramp down magnet per proper GE ramp down procedure for specific magnet
Physical de-install
Verify system Coil List for system, extra Equipment list
Introduce – Talk to Radiology Manager/ Director; Get Emergency /Contact Number; After hours access
Radiology Manager
Phone Number
After Hours Access Key
After Hours Access Code
Get Tools and Supplies
Ramp down kit, Cables, probes, hold down tool
Uline shipment; pallets, bubble wrap, shrink wrap
Identify Magnet exit path, any problems, steel beams, window cracks, height, width, corners, and or soft ground
Verify with Rigger
Identify Equipment exit path, doorways, raised floor, loading dock, pallet jack access, staging area
Chiller locations
De-Cable and Package
Saving Magnet (leave Helium PPJ10 and Pressure PPJ48 cables connected)
Disconnect cables and equiptment
Separate out any equipment and cables customer is keeping (injectors, stereos, ect....)
Disconnect all cables from upper part of pen pannel, both sides (watch for coldhead power on bottom, if it is running)
Disconnect all cables from Cabinets (Systems, RF, Gradient, Gram, PDU, or Variations.....)
Disconnect all cables from Console area
Disconnect all cables from Magnet Room (cover inside of bore and put gradient cables and hose inside)
Leave Magnet Monitor or helium meter connected when possible (if in systems cabinet remove and put with magnet)
Packaging
Wrap cabinet corners with large bubble wrap and shrink wrap over them (remove keys and tape to back of cabinets)
Try to keep cables separate in boxes
One box pallet for equipment room cables and spare parts
One box pallet for magnet room cables and spare parts
One small box console room cables and spare parts
Rig out
Two Hours before Riggers /Crane - Turn off Coldhead, De-install Lines, Compressor and Main Disconnect
Helium Level:
Magnet Pressure
De-install coldhead Lines, Compressor and Main Disconnect
Supervise Equipment Removal, Magnet Rig Out.
Verify all pallets and equipment are in Shipping Container.
Verify Magnet Secure and S/C Lead Extension is Dis-Engaged
Magnet Pressure
Sign off Rigger
Sign off Shipper
Delivery Location
Delivery Arrival Date
Delivery Arrival Time
Street Adress
City
State
Zip Code
Final Report
Final Check List
Items list with stickers
Everything packaged and wrapped to go with the system
Make sure everything gets on the truck and secured so nothing will be damaged
Make sure magnet shim lead is up and record the final helium and pressure reading
Make sure customer is happy and everything is cleaned up when we leave
Helium
Pres.
Date
Time
Data Sheets Completed
Ramp down
Cryogen fill
System Inspection Form
De-install final report
Notes of Follow up Items (Verify that everything is completed)
Signatures of Completion

## Refurbish
Walk Through | General and Outside | Electrical | Mechanical | Rear Room | Scan Room | Operators Room | Equipment Room
Mobile condition fornt | Extra keys to front door | Condition of RussellStoll power connector | Clean Coils for AC and Chillers | Ceiling clean and in good condition
Mobile condition drivers side | Extra keys to back door or lock code | Check shore/generator manual transfer switch | Check and adjust fluid levels for AC and Chillers | Walls clean and in good condition
Mobile condition rear | Extra keys to belly boxes and lift gate door | Check all power and ground connections | Check and adjust pressures for AC and Chillers | Floors clean and in good condition
Mobile condition passanger side | Licence plate current tags | Check 12V DC battery and charger/inverter | Verify temperature of equipment room AC and heater | Plate for securing gradient cables on the wall
Exterior paint | registration for mobile | Patient lift and roll door working properly with both remote and panel controls | Verify temperature of operator and scan room AC and heater | Rear door hardware and lock good
Landing gear | Hud sticker on unit | Mobile electrical connector in good condition | Make sure everything is operational and clean on AC and Chillers | Rear door seals and RF mesh good | Optional equipment mounted and secure, Injector system, stereo system, camera, etc. | Main entry door and Emergency exit door hardware working (exit signs and path marked) | Floor clean and in good condition, cabinets mounted, phantoms in holders with nothing else on floors
Support stands | Twin I-beams and frame in good condition | All break and turn signal covers on and operational | Patient lift hydrolics and motors checked | Top and or side hatch hardware good | Optional equipment all working and operational with everything included with them | Floor clean and in good condition, computers or cabinets mounted with nothing else on floors | All equipment and items in cabinets mounted and secured, covers on and doors locked
Twin I beams and frame | Landing gear and suport stands checked out | All running light covers on and working | Patient lift cleaned and greased | Top and or side hatch seals and RF mesh good | Floor clean and in good condition, anything on floors secured and clean | Counter tops clean and in good condition, everything on them mounted or secure for moving | Condition of storage cabinets, everything cleaned with only necessary items
Tires, rims and breaks | Tires in good condition and pressure verified | Exterior light for entry operational | Check diesel fuel level and guage | Pen pannel secure with hardware and RF seals good | Magnet covers, table, coils, pads and phantoms all cleaned and stored in a secure location | Condition of storage cabinets, everything cleaned with only necessary items | All lights, outlets and switches working, cover plates on in good condition and no exposed wires
Air bags and air system | All lug nuts and anti-lock breaks checked | All covers and breaker panels labeled and secure | Perform PM on generator and change oil and filters | Vent pipe in place and secure | Ceilings clean and in good condition, smoke alarm or fire detection working | Current save Info, MOD options and manuals, folder with all data sheets and log book in cabinet | Ceilings clean and in good condition, smoke alarm or fire detection working
Exterior lights | Air bags and axils on mobile in good condition | Exaust fan in rear of mobile operational | Verify operation and fix any issues with generator | Patient blower secured down | Conditions of wall coverings, holes patched and walls clean in good condition | All lights, outlets and switches working, cover plates on in good condition and no exposed wires | Conditions of wall coverings, holes patched and walls clean in good condition
AC units | Fire detection system checked out | Lights, switches and outlets in rear of mobile checked/replaced | All covers in place for AC and generator | Ladder secured down | All electrical outlets, light switches and e-stop cover plates in place and clean | Ceilings clean and in good condition, smoke alarm or fire detection working | All cables routed correctly, zip tied, clean and secured in place with service loops where needed
Chiller | Two current non ferrous fire extinguishers installed | Lights, switches and outlets in scan room checked/replaced |  |  | All cables routed correctly, zip tied, clean and secured in place with service loops where needed | Conditions of wall coverings, holes patched and walls clean in good condition
Generator | Roof of mobile clean and not leaking | Lights, switches and outlets in operators room checked/replaced |  |  | Pen panel, patient blower and rear pedistool in place and secured, covers and door in good condition | All cables routed correctly, zip tied, clean and secured in place with service loops where needed
Lift gate | Roof hatch operational and seals | Lights, switches and outlets in equipment room checked/replaced |  |  | Vent pipe in place and secured with vent glass, exhaust fan operational | Optional equipment mounted and secure, Injector system, stereo system, camera, etc.
Stairs | Rear or side opening sealed and compleat | Non farrious smoke detectors up to date in each room |  |  | All seals and RF mesh in place | Optional equipment all working and operational with everything included with them
Railings | Vent pipe exit screen and deflector intacked
Front door | Exterior paint and body free of dents, rust or scratches
Rear door | Exterior welding, rivets and bolts sceure with no holes
Belly doors | Top rain guard and corners look good
Roll up or sliding door | Belly doors check for corrosion and leakage
Scan room door | Belly door locks all working
Equipment room door | Belly boxes cleaned out with only necessary items; leveling wood, phone, network, power cord ect…
Top or side hatch | Awning over lift gate operational and clean
Outside roof of mobile | Awning handle and emergency wrench in storage place
Ceiling scan room | Rain deflector at entry door in place and working
Ceiling operators room | Lift gate railings secure with all hardware and working properly
Ceiling equipment room | Stairs stable with rails and all bolts or hardware
Ceiling rear room
Lights scan room
Lights operators room
Lights equipment room
Lights rear room
Walls scan room
Walls operators room
Walls equipment room
Walls rear room
Cabinets scan room
Cabinets operators room
Cabinets equipment room
Countertops scan room
Countertops operators room
Sink
Flooring scan room
Flooring operators room
Flooring equipment room
Flooring rear room

## Install
Install Template
Project ID | Job | MRI Tasks |  |  |  | Job | CT Tasks |  |  |  | Job | PET Tasks
Project Name |  |  | Estimated hrs | Actual hrs | Initals |  |  | Estimated hrs | Actual hrs | Initals |  |  | Estimated hrs | Actual hrs | Initals
| Pre-site Inspection |  |  |  |  | Pre-site Inspection |  |  |  |  | Pre-site Inspection
|  | Verify room dimentions |  |  |  |  | Schedule inspection
|  | Magnet and equipment entrance path |  |  |  |  | Verify equipment and Gantry exit paths
|  | Equipment location and clearance |  |  |  |  | Verify system operation, tube hrs, options and any errors
|  | Magnet mounting and seismic |  |  |  |  | Schedule rig out date
480V power and chiller
| Staging |  |  |  |  | Staging |  |  |  |  | Staging
Clean and paint any damaged covers
Clean and paint table and equipment
Clean filters and cabinets
Re-package system to be delivered to site
| Site readiness | Verify site power connected to main disconnect panel |  |  |  | Site readiness |  |  |  |  | Site readiness
Verify site water and or chiller running with hose barb connections
Verify with contractor opening for magnet and equipment
Verify scheduled rigging company ready and outline of job
| Shipping |  |  |  |  | Shipping |  |  |  |  | Shipping
|  | Stage and verify all equipment is clean ready and packed |  |  |  |  | System cleaned up and staged for shipping
|  | Load equipment and magnet on truck |  |  |  |  | All dollies and moving equipment with system
|  | Sign off that everything is on and secured |  |  |  |  | Load system and equipment on truck
| System Delivery |  |  |  |  | System Delivery | Verify fork Lift and or tow truck for Gantry and equipment |  |  |  | System Delivery
|  | Move Magnet into place and mount down |  |  |  |  | Verify packing supplies and or U-line Shipment
|  | Hook up coldhead compresor |  |  |  |  | Verify everything on truck and sign off
Install vent pipe
Verify helium level and pressure
Check for any damage or missing componets
Unload all equipment from truck and sign off
Move all equipment into place
Install pen panel
Install Magnet monitor and ERU
Verify cable lengths
Check for any damage or missing componets
| Physical Installation |  |  |  |  | Physical Installation |  |  |  |  | Physical Installation
Verify contractors completed and or out of the way
Install cables in magnet room
Install cables in equipment room
Install cables in console room
Verify cable map and connections
List and order any needed cables or parts
| Power Up |  |  |  |  | Power Up |  |  |  |  | Power Up
Verify all power and ground connections
Make sure all breakers are turned off to equipment and the main.
Measure input power before turning on each breaker
Verify correct phasing of equipment
Fix any error or issues to get the system powered up all the way
Run head and body noise scans to check ROIs
Install all covers and dress cables
Verify table and bridge level and operational
Verify patient alert, belows, EKG operational
Verify Bore lights and patient fan operational
Verify compressor and gradient chiller fluid level and temp
Verify emergency stops and emergency offs
Verify magnet signs and gauss line excusions
Verify network setup and configurations
Verify optional equipment; stereo, injector, ect…
List and order any needed coils, phantoms, pads, software, books or parts
Verify LBK currents and Magnet Serial #
Verify that cryogen are ordered and when to start calibrations
Record Magnet and compressor Readings
| Calibrations |  |  |  |  | Calibrations |  |  |  |  | Calibrations
Verify helium level and delivery of product and or tools
De-ice and verify connections
Check for metal and place magnet signs
Ramp magent following procedure
Put in LBK shim currents
Run first image and forward reverse quaduture
Geometric verification auto l-coil and DC offsets
Rough shim until everything is under 50hz
Passive shim magnet (if needed)
Iso center and gradient cals (if twin speed run zoom mode as well)
Run grafidy (if twin speed run zoom mode as well)
Re-run L-coil, Dc offsets
Re-run Iso center and gradient cals (if twin speed run zoom mode as well)
Re-run main supercon shim
Run gradient shim (if twin speed run zoom mode as well)
Verify center frequency and re-park if needed
Check system for ghosting and artifacts
Run DWI white noise checks in all directions
Run system gain cals for head and body
Run Echo Planner Imaging scans for head and body (EPT)
Final iso center and gradient cals
Check multicoil scans on all recevers
Run SNR on all coils
Setup network information including PACS, worklists and camera
Make a save INFO, save GI and protocols
Copy all data sheets for site and office
Clean up everything and get ready to turn over to customer
| System Final Installation |  |  |  |  | System Final Installation |  |  |  |  | System Final Installation
Make a list of all coils on site and check against the quote of coils sold
Verify system operation, system boots up without errors or problems, everything is configured properly
Verify all table functionality; table undock/dock, up/down, drive in/out, and dog house re-tracks
Verify all table pads, coil pads, positioning pads, patient straps and phantoms with system and put away
Verify all extra options and equipment installed; injector, camera, stereo, ect...
Verify network Connections, worklist and camera to make sure they are working
Verify User manuals, coils books, MOD options with system and current save info done
Verify Communication log and all data sheets left in a book at the site with a copy for the office
Verify everything is cleaned up and put away
Make a list of anything that is missing or needs follow up
Final list of all parts used on system
Verify patient scan with customer
Verify one last head, body and phased array scan before leaving
Verify all data sheets are completed; Ramp, Shim, Grad shim, Iso, SNR, Final Report
Final walk through to check that everything is done and clean with all covers in place
Turn system over to the customer and have them sign off that everything is completed

## Finalization
Finalization
Verify PMs done on Unit and Chillers
Secure Keyboard, monitor and mouse
Bungee magnet covers
Lower table, lock wheels and make sure it is secure
Check license plate, tags, registration and HUD are on vehicle
Check to make sure shim lead is up
Check Helium Level and Pressure
Lift gate painted and clean with caution markings

## Old Data
Mobile Template
Project ID | Job | MRI Mobile Tasks | Estimated hrs | Actual hrs | Initals | CT Mobile Tasks | Estimated hrs | Actual hrs | Initals | PET Mobile Tasks | Estimated hrs | Actual hrs | Initals
Project Name
Walk Through
Exterior Paint
Lift Gate
Stairs
Railings
Front Door
Rear Door
Belly Doors
Roll up or Sliding Door
Magnet Room Door
Equipment Room Door
Top Hatch
Outside roof of mobile
Ceiling Magnet room
Ceiling Operators room
Ceiling Equipment room
Lights Magnet room
Lights Operators room
Lights Equipment room
Walls Magnet room
Walls Operators room
Walls Equipment room
Cabinets Magnet room
Cabinets Operators room
Cabinets Equipment room
Countertops Magnet room
Countertops Operators room
Sink
Flooring Magnet room
Flooring Operators room
Flooring Equipment room
Extra Keys to front door, Back door, belly boxes, and lift gate door
Licence Plate current tags and registration for mobile
Hud sticker on unit
Twin I-beam, landing legs, and support stands on chassis (check for cracks or damage)
Verify tire tread, wear, and pressure check that lug nuts are all there, Anti-lock brakes (if available)
Condition of roof, roof hatch and magnet exit
Electrical
Condition of RussellStoll power connector, cable secure and lugs tight
Check all power and ground connections, all covers and breaker panels labeled and secure
Check 12V DC battery and charger/inverter
AC and Chillers
Clean Coils
Check and adjust fluid levels
Generator
Previous check sheet
Condition of RussellStoll power connector, cable secure and lugs tight
Check all power and ground connections, all covers and breaker panels labeled and secure
Check 12V DC battery and charger/inverter with 12V circuit breaker panel and master switch
|  | , all seals and RF mesh in place |  |  |  | Working condition of HVAC units, coils cleaned and system PM complete everything in spec
|  | Working condition of HVAC units, coils cleaned and system PM complete everything in spec |  |  |  | Check working condition of A/C and heater in equipment, operator and magnet rooms
|  | Check working condition of A/C and heater in equipment, operator and magnet rooms |  |  |  | Twin I-beams, landing legs, and support stands on chassis (check for cracks or damage)
Verify tire tread, wear, and pressure check that lug nuts are all there, Anti-lock brakes (if available)
Condition of roof, roof hatch and gantry exit, all seals and Lead lining in place
Fire detection system and current fire extinguisher installed in at least 2 locations
|  | Fire detection system and current non ferrous fire extinguisher installed in at least 2 locations |  |  |  | Optional equipment mounted and secure, Injector system, stereo system, camera, etc.
|  | Optional equipment mounted and secure, Injector system, stereo system, camera, etc. |  |  |  | Optional equipment all working and operational with everything included with them
|  | Optional equipment all working and operational with everything included with them |  |  |  | Check diesel fuel level and generator operation, verify condition
|  | Check diesel fuel level and generator operation, verify condition |  |  |  | Check shore/generator manual transfer switch
Check store/generator manual transfer switch
| Exterior Condition Front |  |  |  |  | Paint and body clean, no dents, rust or scratches and in good condition
|  | Paint and body clean, no dents, rust or scratches and in good condition |  |  |  | Welding, riveting and bolts on exterior all in place no holes, all covers in place for AC and generator
|  | Welding, riveting and bolts on exterior all in place no holes, all covers in place for AC and generator |  |  |  | Air break and electrical connectors in good condition, all running lights and covers on
Air break and electrical connectors in good condition, all running lights and covers on
| Exterior Condition Street Side |  |  |  |  | Paint and body clean, no dents, rust or scratches and in good condition
|  | Paint and body clean, no dents, rust or scratches and in good condition |  |  |  | Welding, riveting and bolts on exterior all in place no holes, top rain guard and corners look good
|  | Welding, riveting and bolts on exterior all in place no holes, top rain guard and corners look good |  |  |  | Running lights and turn signal covers on and working, leveling feet and pads all there and in place
|  | Running lights and turn signal covers on and working, leveling feet and pads all there and in place |  |  |  | Lower storage doors check for corrosion, leakage and door locks working, clean and look good
|  | Lower storage doors check for corrosion, leakage and door locks working, clean and look good |  |  |  | Lower storage bays cleaned out, only necessary items; leveling wood, phone, network and power cords
Lower storage bays cleaned out, only necessary items; leveling wood, phone, network and power cords
| Exterior Condition Rear |  |  |  |  | Paint and body clean, no dents, rust or scratches and in good condition
|  | Paint and body clean, no dents, rust or scratches and in good condition, vent pipe exit screen and deflector |  |  |  | Welding, riveting and bolts on exterior all in place, rear door hardware, lock, and lead good and clean
|  | Welding, riveting and bolts on exterior all in place, rear door hardware, lock, and RF seal good and clean |  |  |  | All break and running light covers on and operational, license plate in place with current tags
All break and running light covers on and operational, license plate in place with current tags
| Exterior Condition Curb Side |  |  |  |  | Paint and body clean, no dents, rust or scratches and in good condition
|  | Paint and body clean, no dents, rust or scratches and in good condition |  |  |  | Welding, riveting and bolts on exterior all in place no holes, top rain guard and corners look good
|  | Welding, riveting and bolts on exterior all in place no holes, top rain guard and corners look good |  |  |  | Running lights and turn signal covers on and working, leveling feet and pads all there and in place
|  | Running lights and turn signal covers on and working, leveling feet and pads all there and in place |  |  |  | Lower storage doors check for corrosion, leakage and door locks working, clean and look good
|  | Lower storage doors check for corrosion, leakage and door locks working, clean and look good |  |  |  | Lower storage bays cleaned out, only necessary items; leveling wood, phone, network and power cords
|  | Lower storage bays cleaned out, only necessary items; leveling wood, phone, network and power cords |  |  |  | Awning over lift gate operational and clean, awning handle and emergency wrench in storage place
|  | Awning over lift gate operational and clean, awning handle and emergency wrench in storage place |  |  |  | Patient lift and roll door working properly with both remote and panel controls
|  | Patient lit and roll door working properly with both remote and panel controls |  |  |  | Lift gate railings secure with all hardware and working properly, everything clean with caution markings
|  | Lift gate railings secure with all hardware and working properly, everything clean with caution markings |  |  |  | Aluminum stair assembly stable with rails and all bolts or hardware
|  | Aluminum stair assembly stable with rails and all bolts or hardware |  |  |  | Main entry door and Emergency exit door hardware working (exit signs and path marked)
|  | Main entry door and Emergency exit door hardware working (exit signs and path marked) |  |  |  | Exterior light and rain deflector drip rail at entry door in place and working
Exterior light and rain deflector drip rail at entry door in place and working
| Operators Room |  |  |  |  | Floor clean and in good condition, computers or cabinets mounted with nothing else on floors
|  | Floor clean and in good condition, computers or cabinets mounted with nothing else on floors |  |  |  | Counter tops clean and in good condition, everything on them mounted or secure for moving
|  | Counter tops clean and in good condition, everything on them mounted or secure for moving |  |  |  | Condition of storage cabinets, everything cleaned with only necessary items
|  | Condition of storage cabinets, everything cleaned with only necessary items |  |  |  | Current save state, MOD options and manuals, folder with all data sheets and log book in cabinet
|  | Current save Info, MOD options and manuals, folder with all data sheets and log book in cabinet |  |  |  | Ceilings clean and in good condition, smoke alarm or fire detection working
|  | All lights, outlets and switches working, cover plates on in good condition and no exposed wires |  |  |  | Conditions of wall coverings, holes patched and walls clean in good condition
|  | Ceilings clean and in good condition, smoke alarm or fire detection working |  |  |  | All cables routed correctly, zip tied, clean and secured in place with service loops where needed
Conditions of wall coverings, holes patched and walls clean in good condition
|  | All cables routed correctly, zip tied, clean and secured in place with service loops where needed |  |  |  | Floor clean and in good condition, cabinets mounted, phantoms in holders with nothing else on floors
| Equipment Room |  |  |  |  | All equipment and items in cabinets mounted and secured, covers on and doors locked
|  | Floor clean and in good condition, cabinets mounted, phantoms in holders with nothing else on floors |  |  |  | Condition of storage cabinets, everything cleaned with only necessary items
|  | All equipment and items in cabinets mounted and secured, covers on and doors locked |  |  |  | All lights, outlets and switches working, cover plates on in good condition and no exposed wires
|  | Condition of storage cabinets, everything cleaned with only necessary items |  |  |  | Ceilings clean and in good condition, smoke alarm or fire detection working
|  | All lights, outlets and switches working, cover plates on in good condition and no exposed wires |  |  |  | Conditions of wall coverings, holes patched and walls clean in good condition
|  | Ceilings clean and in good condition, smoke alarm or fire detection working |  |  |  | All cables routed correctly, zip tied, clean and secured in place with service loops where needed
Conditions of wall coverings, holes patched and walls clean in good condition
|  | All cables routed correctly, zip tied, clean and secured in place with service loops where needed |  |  |  | Floor clean and in good condition, anything on floors secured and clean
| Magnet Room |  |  |  |  | Magnet covers, table, coils, pads and phantoms all cleaned and stored in a secure location
|  | Floor clean and in good condition, anything on floors secured and clean |  |  |  | Ceilings clean and in good condition, smoke alarm or fire detection working
|  | Magnet covers, table, coils, pads and phantoms all cleaned and stored in a secure location |  |  |  | Conditions of wall coverings, holes patched and walls clean in good condition
|  | Ceilings clean and in good condition, smoke alarm or fire detection working |  |  |  | All electrical outlets, light switches and e-stop cover plates in place and clean
|  | Conditions of wall coverings, holes patched and walls clean in good condition |  |  |  | All cables routed correctly, zip tied, clean and secured in place with service loops where needed
All electrical outlets, light switches and e-stop cover plates in place and clean
|  | All cables routed correctly, zip tied, clean and secured in place with service loops where needed |  |  |  | Verify PMs done on Unit and Chillers
|  | Pen panel, patient blower and rear pedistool in place and secured, covers and door in good condition |  |  |  | Secure Keyboard, monitor and mouse put away and anything on counters
|  | Vent pipe in place and secured with vent glass, exhaust fan operational |  |  |  | Secure Gantry, table phantoms and covers
| Final Checks Before Moving Mobile |  |  |  |  | Check license plate, tags, registration and HUD are on vehicle
Verify PMs done on Unit and Chillers
Secure Keyboard, monitor and mouse
Bungee magnet covers
Lower table, lock wheels and make sure it is secure
Check license plate, tags, registration and HUD are on vehicle
Check to make sure shim lead is up
Check Helium Level and Pressure