GE PM 9x – 12x

TAKE PICTURES OF EVERYTHING

- Arrive on site – talk to customer / tech and see what issues they have been having if any

- Check previous scans for any issues / see what coils they use the most

- Do a save info to check the MOD / CD drives and make sure they are functioning

- Set date and time (after restart, login to root – operator to set date and time)

- Go to scan room – test patient alert systems (squeeze ball, Gating Cable/ Respiratory Bellows / PG Probe)

- Check the patient table release (at the end of the table) and the arm boards (sides of table)

- Check the dock emergency release – pull out handle and push it back in – redock table

- Run the table halfway in – press on the down foot lever to ensure the table does not go down when in the bore

- Unhook the patient table – flip it over and check each wheel to ensure they spin (once a year, switch out or soak the wheels)

- Leave table at the end – run the rest all the way to the back of the magnet

- Inspect the bore for any metal

- Go to the doghouse (back of the magnet) check the cable release lever (under back of doghouse)

- Flip over the table motor – check the wheels

- Remove doghouse side covers – check underneath for any metal with a flashlight

- Set table back to how it was – Run it to home

- Check the coldhead for any condensation

- Clean the patient table, and front covers with a magic eraser

- Put the head coil and DQA phantom on the table

- Go to the host computer – shut down the system (press red button on cabinets when down)

- Vacuum all the filters on the cabinets

- Write down all data on the magmon – get compressor hours, pressure, and check water lines / helium lines

- Inspect chiller – ensure there is enough water / glycol in it

- Return to the console area – vacuum out the console fans / graphics card

- Turn on cabinets (pull the lever up, push it down, hit green button)

- Twice a year – check ERU, check all cables on pin panel to ensure they are tight, check magnet room door frame, check water chiller filter, check the placement of magnet room warning signs, test emergency stop buttons (NOT the quench button)

- Once system is booted up, start the test scans

- Run ISO center (see word document for steps to run ISO center)

- TPS reset after ISO (9x must shut down system)

- Run DQA calibration (see word document for steps on DQA)

  - Run grad cal (Cals / check – NEMA)

    - Two images for X, Y, and Z

    - Ensure in DQA cal the correct images are selected

    - Write down new and old numbers – find the difference

    - Go to research options, display CVs – (enter what it says in top blank – example cfxfull)

    - WRITE DOWN OLD NUMBER

    - DO NOT WRITE IN NEW NUMBER – add or subtract the difference of the two numbers to the number shown in CVs

- Check Head SNR (TLT phantom in head coil) (see word document for steps, scan on axial, sagittal and coronal)

- Check Body SNR (Big TLT phantom, no coil) (see word document for steps, scan on axial, sagittal, and coronal)

- Scan with CTL coil – two phantoms, set up first scan on top portion

  - New patient

  - Patient ID – geservice

  - Patient name – ctl test

  - Patient position – supine

  - Patient entry – head first

  - Coil – ctl top (then ctl bottom)

  - Plane – sagittal

  - Mode – 2D

  - Pulse seq – spin echo

  - Grad mode – whole (zoom as well if fucntioning

  - # of TEs per scan – 1

  - TE – minimum

  - TR – 300

  - Bandwidth – 15.63

  - FOV – 48

  - Slice thickness – 3

  - Spacing – 0

  - Frequency – 256

  - Phase – 256

  - NEX – 1

  - Phase FOV – 1

  - Freq DIR – S/I

  - CF – peak

  - R / L start – 0

  - R / L end – 0

  - # of slices – 1

- Save series – auto prescan, manual prescan

- Before scan – right click research options, display CVs , Saveinter, current value change to 1

- Scan – check images

- Change landmark to lower portion – scan ctl bottom

- Scan one other coil that is used (scan all coils if there aren’t that many)

- Run DC offsets – no coil needed

- TPS reset after DC offsets

- Run LV shim (dino ball) – see word document for LVshim steps

- Once scans are complete – C shell, storelog, enter

- Save info – guided install, save / restore, save info

- Remove all used phantoms and coils

- Follow the sites shut down procedures (if they leave system up or shut it down at night)

- Update site logbook, clean up