Site:				               S/N:				 	Date:

Contract Type:

Performed by:

|  | Helium Level % | CCU | Cond | Flow 1 | Flow 2 | Temp 1 | Temp 2 |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Reading |  |  |  |  |  |  |  |
| Alarm set point | N/A | N/A | 7.0 | 3.5 | 3.5 | 28 | 28 |

|  | Pass | Fail |  | Flow (GPM) | Water Temp (to compressor) | Pressure (to compressor) | Pressure (to sense) | Coldhead S/N |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Water Visible? (in Sight Glass) |  |  |  |  |  |  |  |  |
| Disk Drive |  |  |  |  |  |  |  |  |
| Alignment Lights |  |  |  | Compressor Hours | Supply Pressure | Center Frequency |  |  |
| Autodialer Check |  |  |  |  |  |  |  |  |

| Head | GAIN | H1 | SNR | J/F M/A M/J J/A S/O N/D N/A |
| --- | --- | --- | --- | --- |
| Actual |  |  |  |  |
| Specification | N/A | <330 | ≥140 |  |

| Knee | GAIN | H1 | SNR | J/F M/A M/J J/A S/O N/D N/A |
| --- | --- | --- | --- | --- |
| Actual |  |  |  |  |
| Specification | N/A | N/A | ≥310 |  |

| Quad C-Spine | GAIN | H1 | SNR | J/F M/A M/J J/A S/O N/D N/A |
| --- | --- | --- | --- | --- |
| Actual |  |  |  |  |
| Specification | N/A | N/A | ≥145 |  |

| Wrist | GAIN | H1 | SNR | M/A S/O N/A |
| --- | --- | --- | --- | --- |
| Actual |  |  |  |  |
| Specification | N/A | N/A | ≥290 |  |

| Joint/Neck | GAIN | H1 | SNR | M/J N/D N/A |
| --- | --- | --- | --- | --- |
| Actual |  |  |  |  |
| Specification | N/A | N/A | ≥120 |  |

| LG Extremity | GAIN | H1 | SNR | M/J N/D N/A |
| --- | --- | --- | --- | --- |
| Actual |  |  |  |  |
| Specification | N/A | N/A | ≥120 |  |

| Med Flex | GAIN | H1 | SNR | J/F M/A M/J J/A S/O N/D N/A |
| --- | --- | --- | --- | --- |
| Actual |  |  |  |  |
| Specification | N/A | N/A | ≥85 |  |

| Large Flex | GAIN | H1 | SNR | J/F M/A M/J J/A S/O N/D N/A |
| --- | --- | --- | --- | --- |
| Actual |  |  |  |  |
| Specification | N/A | N/A | ≥100 |  |

| XL Flex | GAIN | H1 | SNR | M/A S/O N/A |
| --- | --- | --- | --- | --- |
| Actual |  |  |  |  |
| Specification | N/A | N/A | ≥85 |  |

| MA Shoulder | GAIN | H1 | SNR (1) | SNR (2) | M/A S/O N/A |
| --- | --- | --- | --- | --- | --- |
| Actual |  |  |  |  |  |
| Specification | N/A | N/A | ≥75 | ≥160 |  |

CTL

| Mode (R Coil) | Position | SNR | Specification | M/J N/D N/A |
| --- | --- | --- | --- | --- |
| C (A) | 20,0 |  | ≥190 |  |
| CT (B) | 25,150 |  | ≥80 |  |
| CT (B) | 25,0 |  | ≥89 |  |
| CT (B) | 25,-100 |  | ≥92 |  |
| T (B) | 25,100 |  | ≥82 |  |
| T (B) | 25,0 |  | ≥95 |  |
| T (B) | 25,-100 |  | ≥81 |  |
| L (B) | 25,80 |  | ≥86 |  |
| L (B) | 25,-20 |  | ≥125 |  |
| L (B) | 25,-80 |  | ≥125 |  |

Protocol backup:                     	 Jan/Feb      Jul/Aug	 N/A

Service backup:                        	 Jan/Feb      Jul/Aug	 N/A

Change DI Filter:                    	 Jan/Feb	 N/A

Change particulate filter:       	 Jan/Feb 	 N/A

Change O2 sensor:                   	 Jan/Feb	 N/A

Clean RF room door:               	 May/Jun	 N/A

Check/Adjust FOV’s:		 Jul/Aug	 N/A

Adjust Characteristic Matrix:	 Sep/Oct	 N/A

Clean All Air Filters:		 May/Jun	 Nov/Dec	 N/A

Notes: