# Compressor Test

- What is the incoming power on a Sumitomo water cooled compressor?

480 volts

- How long is a Adsorber good for?

27,000 hours

- What should the water temps be?

10 C to 20 C

- What should the water flow be?

1.3 gpm water to 2.6 gpm glycol

- Approximately what should the static and dynamic pressure be on CSW’s?

Static – 1.6 to 1.65 MPa

Dynamic – 2.1 to 2.3 MPa

- How Should the gas lines feel?

Supply line should be cold

Return line should be warm

- What is oil contamination?

Visible by black light on the connection on the cold head / compressor lines