## Magnet Info
OEM | Magnet Models | Field Strength | Serial Number Identifier | System Model | Allowable Boiloff /wk | Helium Capacity | Verified capacity | OEM Burnoff /yr | AP Burnoff /yr | Pressure range | Min Fill Level | Min Ramp Level | Min Shim Level | Gain per 250L | Notes
Alltech | Centauri | 1.5T |  | Centauri |  | 2200 | [7 empty] | 11%
Alltech | Echostar | 1.5T |  | Echostar
Elscint | 920 | 2.0T |  | Prestige |  | 780 | [7 empty] | 32%
Elscint | Gyrex | 0.5T |  | Gyrex |  | 780 | [7 empty] | 32%
Elscint | OTB | 2.0T |  | OTB |  | 1000 | [7 empty] | 25%
Elscint | Privilege | 0.5T |  | Privilege |  | 330 | [7 empty] | 75%
GE | 3T94 | 3.0T |  | 11.x - 12.x |  | 3066 |  |  |  | 0.25-0.5 psi |  |  |  | 8%
GE | 450w | 1.5T |  | >=23.x | 0.5% | 1769 |  | 0 | 0 | 2psi |  |  |  | 14% | No actual boiloff, 0.5% weekly value is just to allow for fluctuation.
GE | 7.0T | 7.0T |  |  | 0.5% | 1323 |  | 0 | 0 | 2psi |  |  |  | 18% | No actual boiloff, 0.5% weekly value is just to allow for fluctuation.
GE | 750w | 3.0T |  | >=23.x | 0.5% | 1818 |  | 0 | 0 | 2psi |  |  |  | 13% | No actual boiloff, 0.5% weekly value is just to allow for fluctuation.
GE | 9.0T | 9.0T |  |  | 0.5% | 2253 |  | 0 | 0 | 2psi |  |  |  | 11% | No actual boiloff, 0.5% weekly value is just to allow for fluctuation.
GE | Cx 1.5T | 1.5T | Nxxxx | <= 12.x | 1.563% | 1940 | Yes | 1577 |  | 0.25-0.5 psi | 42% | 90% | 75% | 12%
GE | Cx 1T | 1.0T | Pxxxx | <= 12.x | 1.539% | 1970 | Yes | 1577 |  | 0.25-0.5 psi | 42% | 90% | 75% | 12%
GE | HFO | 0.7T |  | OpenSpeed | 0.5% | 376 | Yes | 0 | 0 | 0.9-1.1 psi or <5.0 psi | 20% T 50% B | 95% T 100% B | NA | 150% | No actual boiloff, 0.5% weekly value is just to allow for fluctuation. Capacity is 170L top, 170L bottom, 36L in posts.
GE | LCC 1.5T | 1.5T | Rxxxx (RA also) | 9.x - current | 0.5% | 1940 | Yes | 0 | 0 | 3.0-5.0 psi | 50% | 90% | 75% | 12% | No actual boiloff, 0.5% weekly value is just to allow for fluctuation. RAxxxx serial number indicates experimental winding (require different ramp profile, only 26 exist).
GE | LCC 1T | 1.0T | Qxxxx (QR also) | 9.x - 12.x | 0.5% | 1970 | Yes | 0 | 0 | 3.0-5.0 psi | 50% | 90% | 75% | 12% | No actual boiloff, 0.5% weekly value is just to allow for fluctuation. If QRxxxx the Q is typically hand stamped. These were originally 1.5T magnets that were found to only ramp to 1.0T in Florence.
GE | LCC300 | 3.0T |  | 11.x - current | 0.5% | 1735 | Yes | 0 | 0 | 3.0-5.0 psi | 50% | 90% | 75% | 14% | No actual boiloff, 0.5% weekly value is just to allow for fluctuation.
GE | MAX | 0.5T |  | <= 12.x | 2.466% | 1366 | Yes | 1752 |  | 0.25-0.5 psi | 50% | 75% | 50% | 18%
GE | ONI 1.5T | 1.5T |  | MSK 1.5 | 0.5% | 50 |  | 0 | 0 |  |  |  |  | 500% | No actual boiloff, 0.5% weekly value is just to allow for fluctuation.
GE | ONI 1T | 1.0T |  | OrthOne | 0.5% | 25 |  | 0 | 0 |  |  |  |  | 1000% | No actual boiloff, 0.5% weekly value is just to allow for fluctuation.
GE | Ovation | 0.35T |  | Ovation | ----- | ----- | NA | ----- | ----- | ----- | ----- | ----- | ----- | ----- | Permanent Magnet
GE | Profile | 0.2T |  | Profile | ----- | ----- | NA | ----- | ----- | ----- | ----- | ----- | ----- | ----- | Permanent Magnet
GE | SI | 1.5T | 67xxxx | <= 9.x |  | 1872 |  |  |  | 0.25-0.5 psi |  |  |  | 13%
GE | SII | 1.5T | Cxxxx | <= 9.x | 2.771% | 1216 | Yes | 1752 |  | 0.25-0.5 psi | 50% | 75% | 50% | 20%
GE | SIII | 1.5T | Dxxxx | <= 12.x | 2.836% | 1188 | Yes | 1752 | 1900 | 0.25-0.5 psi | 50% | 75% | 50% | 21%
GE | SIV | 1.5T | Gxxxx | <= 12.x | 1.75% | 1925 | Yes | 1752 | 3000 | 0.25-0.5 psi | 65% | 90% | 70% | 12%
GE | SV | 1.5T | Jxxxx | <= 12.x | 1.922% | 1753 | Yes | 1752 | 2700 | 0.25-0.5 psi | 50% | 75% | 50% | 14%
GE | SX | 1.0T | Xxxxx (K or L also) | <= 12.x | 1.801% | 1871 | Yes | 1752 | 2800 | 0.25-0.5 psi | 50% | 80% | 50% | 13% | Kxxxx serial number for Xc_1 version. Lxxxx serial number for Xc_2 version.
GE | Vectra | 0.5T |  |  |  | 565 | [7 empty] | 44%
GE | VMX2 | 0.5T |  | Contour |  | 722 |  |  |  | 0.25-0.5 psi |  |  |  | 34%
Hitachi | Airis 1 | 0.3T | Axxxx | Airis 1 | ----- | ----- | NA | ----- | ----- | ----- | ----- | ----- | ----- | ----- | Permanent Magnet
Hitachi | Airis Elite | 0.3T | Hxxxx | Airis Elite | ----- | ----- | NA | ----- | ----- | ----- | ----- | ----- | ----- | ----- | Permanent Magnet
Hitachi | Airis II | 0.3T | Cxxxx | Airis II | ----- | ----- | NA | ----- | ----- | ----- | ----- | ----- | ----- | ----- | Permanent Magnet
Hitachi | Altaire | 0.7T | Mxxxx | Altaire | 0.5% | 700 |  | 0 | 0 |  |  |  |  | 35% | 3 Stage Coldhead! No actual boiloff, 0.5% weekly value is just to allow for fluctuation.
Hitachi | Aperto | 0.4T |  | Aperto | ----- | ----- | NA | ----- | ----- | ----- | ----- | ----- | ----- | ----- | Permanent Magnet
Hitachi | Echelon | 1.5T |  | Echelon | 0.5% |  |  | 0 | 0 | [5 empty] | No actual boiloff, 0.5% weekly value is just to allow for fluctuation.
Hitachi | Oasis | 1.2T |  | Oasis
Philips | Asset | 0.5T |  | Asset |  | 740 | [7 empty] | 33%
Philips | F2000 | 1.5T |  | Achieva/Intera | 0.3% | 1730 |  |  | 300 | 60mb |  |  |  | 14%
Philips | Flint | 1.5T |  | Gyroscan | 3% | 1220 |  |  | 1903 |  |  |  |  | 20%
Philips | Gecko | 1.5T |  | Prodiva |  | 1325 | [7 empty] | 18%
Philips | Ingenia 1.5T | 1.5T |  | Ingenia 1.5T | 0.5% | 1840 |  | 0 | 0 | 40mb (0.6psi) |  |  |  | 13% | No actual boiloff, 0.5% weekly value is just to allow for fluctuation.
Philips | Ingenia 3T | 3.0T |  | Ingenia 3T | 0.5% | 1840 |  | 0 | 0 | 40mb (0.6psi) |  |  |  | 13% | No actual boiloff, 0.5% weekly value is just to allow for fluctuation.
Philips | Panorama .23T | 0.23T |  | Panorama | ----- | ----- | NA | ----- | ----- | ----- | ----- | ----- | ----- | ----- | Permanent Magnet
Philips | Panorama 0.6T | 0.6T |  | Panorama |  | 380 |  |  |  | 40mb (0.6psi) |  |  |  | 65%
Philips | Panorama 1T | 1.0T |  | Panorama |  | 750 |  |  |  | 40mb (0.6psi) |  |  |  | 33%
Philips | Rex | 1.5T |  | Achieva XR |  | 2070 |  |  |  | 40mb (0.6psi) |  |  |  | 12%
Philips | Rex | 3.0T |  | Achieva XR | 0.5% | 2070 |  | 0 | 0 | 40mb (0.6psi) |  |  |  | 12% | No actual boiloff, 0.5% weekly value is just to allow for fluctuation.
Philips | T10 | 1.0T |  | NT |  | 1220/1400/1600 | [7 empty] | 20%/17%/15% | There are three different possible T10 NT magnets.
Philips | T5 | 0.5T |  | NT |  | 1000 | [7 empty] | 25%
Philips | Titan | 3.0T |  | Achieva 3T |  | 1500 |  |  | 2340 | 0.25psi |  |  |  | 16%
Picker | OR26 | 1.0T |  | Polaris |  | 1752 |  |  | 2733 | 0.25psi |  |  |  | 14%
Picker | OR35 | 1.5T |  | Edge |  | 1630 |  |  | 2542 | 0.25psi |  |  |  | 15%
Picker | OR42 | 1.0T | --- | Vista |  | 1328 |  |  | 2072 | 0.25psi |  |  |  | 18%
Picker | OR71 | 1.5T |  | Eclipse |  | 1770 |  |  | 2761 | 0.25psi |  |  |  | 14%
Siemens | C | 0.35T |  | C | ----- | ----- | NA | ----- | ----- | ----- | ----- | ----- | ----- | ----- | Permanent Magnet
Siemens | Concerto | 0.2T |  | Concerto | ----- | ----- | NA | ----- | ----- | ----- | ----- | ----- | ----- | ----- | Permanent Magnet
Siemens | Jazz | 0.2T |  | Jazz | ----- | ----- | NA | ----- | ----- | ----- | ----- | ----- | ----- | ----- | Permanent Magnet
Siemens | Open Viva | 0.2T |  | Open Viva | ----- | ----- | NA | ----- | ----- | ----- | ----- | ----- | ----- | ----- | Permanent Magnet
Siemens | OR103 | 3.0T | --- | Monet/Spectra |  | 1209 |  |  | 1886 | 15.3 psia (0.6 psi) |  |  |  | 20%
Siemens | OR105 | 1.5T | --- | Avanto | 0.5% | 1400 |  | 0 | 0 | 15.3 psia (0.6 psi) | 50% |  |  | 17% | Scanning inhibited below 40% helium, no actual boiloff, 0.5% weekly value is just to account for fluctuation.
Siemens | OR122 | 1.5T | 3 | Espree | 0.5% | 1460 | Yes | 0 | 0 | 15.3 psia (0.6 psi) | 50% |  |  | 17% | Scanning inhibited below 40% helium, no actual boiloff, 0.5% weekly value is just to account for fluctuation.
Siemens | OR20 | 1.0T |  | Rhapsody Open |  | 1140 | [7 empty] | 21%
Siemens | OR24 | 1.0T | --- | Harmony | 2% | 1806 |  |  | 2817 | 15.3 psia (0.6 psi) | 50% |  |  | 13% | 2%/wk estimate
Siemens | OR35 | 1.5T |  | Vision/Sonata |  | 1630 |  |  | 2542 | 15.3 psia (0.6 psi) |  |  |  | 15%
Siemens | OR41 | 1.0T | --- | Impact/Expert |  | 1240 |  |  | 1933 | 15.3 psia (0.6 psi) |  |  |  | 20%
Siemens | OR42 | 1.0T | --- | Impact/Expert |  | 1328 |  |  | 2072 | 15.3 psia (0.6 psi) |  |  |  | 18%
Siemens | OR60 | 1.5T | --- | Essenza |  | 1422 |  |  | 2218 | 15.3 psia (0.6 psi) | 50% |  |  | 17% | Scanning inhibited below 40% helium, no actual boiloff, 0.5% weekly value is just to account for fluctuation.
Siemens | OR64 | 3.0T | --- | Trio | 2.5% | 1200 | Yes |  | 2019 | 15.3 psia (0.6 psi) | 35% |  |  | 20% | 2.5%/wk estimate. Scanning inhibited below 40% helium
Siemens | OR65 | 3.0T | --- | Prisma |  | 1035 |  |  | 1615 | 15.3 psia (0.6 psi) |  |  |  | 24%
Siemens | OR70 | 1.5T | --- | Symphony | 2% | 1304 | Yes | 1356.16 | 2235 | 15.3 psia (0.6 psi) | 50% |  |  | 14% | 2%/wk estimate (can be 1.5-3% depending on magnet age). Poor efficiency fills are common.
Siemens | OR92 | 3.0T | --- | Allegra |  | 691 |  |  |  | 15.3 psia (0.6 psi) |  |  |  | 36%
Siemens | OR93 | 1.5T |  | Amira |  | 1254 |  |  |  | 15.3 psia (0.6 psi) |  |  |  | 19%
Siemens | OR97 | 3.0T | --- | Verio | 0.5% | 1421.0526 |  | 0 | 0 | 15.3 psia (0.6 psi) | 50% |  |  | 17% | Scanning inhibited below 40% helium, no actual boiloff, 0.5% weekly value is just to account for fluctuation. 1350 liters at 95% is maximum fill allowed.
Siemens | OR98 | 1.5T | --- | Aera | 0.5% | 1820 | Yes | 0 | 0 | 15.3 psia (0.6 psi) | 50% |  |  | 13% | Scanning inhibited below 40% helium, no actual boiloff, 0.5% weekly value is just to account for fluctuation.
Siemens | OR99 | 3.0T | --- | Skyra | 0.5% | 1190 |  | 0 | 0 | 15.3 psia (0.6 psi) | 50% |  |  | 21% | Scanning inhibited below 40% helium, no actual boiloff, 0.5% weekly value is just to account for fluctuation.
Toshiba | MRT50 | 0.5T |  | Flexart |  | 459 | [7 empty] | 54%
Toshiba | Opart | 0.35T |  | Opart | ----- | ----- | NA | ----- | ----- | ----- | ----- | ----- | ----- | -----
Toshiba | OR200 | 1.5T |  | Orian |  | 1211 | [7 empty] | 20%
Toshiba | OR76 | 1.5T |  | Vantage (Orian upgrade) | 1% | 1640 | Yes | 2235 |  | 0.25psi |  |  |  | 15%
Toshiba | Titan | 3.0T |  | Titan | 1% | 1500 |  |  | 2340 | 0.25psi |  |  |  | 16%
Toshiba | TN150 | 1.5T |  | Vantage Elan |  | 1255 | [7 empty] | 19%
United |  | 3.0T |  | uMR 790
United |  | 3.0T |  | uMR 780
United |  | 1.5T |  | uMR 570

## Unit Converter
Unit Conversions
Type value you have in the GREEN boxes to get conversions
Pressure
psi (GE) | psia (SM) | mbar (PH) | MPa
1.35 | 16.05 | 93.0798 | 0.0093079
0.6 | 15.3 | 41.3688 | 0.0041369
5.8015 | 20.5015 | 400 | 0.04
29.0074 | 43.7074 | 2000 | 0.2
Flow
Gpm | Lpm
1.5 | 5.6781
1.1095 | 4.2
Lhe% <--> actual liters
Magnet Model | Magnet Capacity | % | Liters
or122 | 1460 | 18.7% | 273.02
OR64 | 1200 | 41.67% | 500 | Nuclei | Factor
|  |  |  | Hydrogen (H) | 4257.6
Magnetic Field |  |  |  | Helium (He) | 3243
Nuclei |  | Hydrogen (H) | 4257.6 | Carbon (C) | 1071
Tesla | Gauss | Frequency (Hz) | Frequency (MHz) | Sodium (Na) | 1126
1.5 | 15000 | 63864000 | 63.864 | Fluorine (F) | 4006
0.0005 | 5 | 21288 | 0.021288 | Phosphorus (P) | 1724
0.3758 | 3757.9857 | 16000000 | 16 | Xenon (Xe) | 1178