## 3.0T HDx Upgrade Checklist
| 3.0T HDx/HDxt Upgrade Quote Checklist | [6 empty] | Current Sys type | Magnet Type | System Type |  | Site Type |  |  | Current Gradient Driver |  |  |  | Channels | Array Processor | [5 empty] | Syscab pn # | New or Upgrade
| JAN 2015, Rev 17 | [26 empty] | 3x | 5127452 | 1 | No MNS
[8 empty] | LX/Eclipse |  | Head only scanner | Roemer | Fixed | Long - 1 Piece | Mag room 19ft, Equip room 25ft | SGD w/SRFD | Date prior to Aug 28, 2001 | Salt Block Horizon | Yes | 1 | Reflex 20 | Front is not SCP 2 | Rear is not Exciter 2 | Rear is not RCVR 2 | Rear is not UTNS-3 | None | 2357500 | New (Forward Production) | 4.0 | 5127452-3 | 2 | 8KW AMT
[8 empty] | Excite II |  | Full body scanner | CRM | Relocatable | Split | Mag room 19ft, Equip room 55ft | ACGD | Date Aug 28, 2002 or later | Wide Open | No | 4 | Reflex 50 | Front is SCP 2 | Rear is Exciter 2 | Rear is RCVR 2 | Rear is UTNS-3 | RRI | 2357500-2 | Upgrade | 4.1 | 5337894-3 | 4 | 4KW CPC
| INSTRUCTIONS | [6 empty] | HD | 3T94 |  | BRM |  |  | Mag room 40ft, Equip room 25ft | HFD or HFD-s |  |  |  | 8 | Reflex 100 |  |  |  |  | Plexus | 5118496-5 |  | 4.1.3 | 5337894-4
[8 empty] | HDx | G3 |  | TRM |  |  | At Excite - No Cables |  |  |  |  | 16 | Reflex 200 (SVP4) | [5 empty] | 5118496-6 |  | 5.2.9 | 5911000-4
| 1). Please fill out all necessary sections (1 - 4) below. | [6 empty] | HDxt | [5 empty] | At Excite HD no cables |  |  |  |  | 32 | Reflex 400 (SVP8) | [5 empty] | 5118496-7 |  | N/A | 5911000-5
[8 empty] | HD23 | [11 empty] | Reflex 800 (SVP16) | [5 empty] | 5110607-2
| 2). If requested, also complete the Upgrade Performance Checklist. | [18 empty] | VRE 2 ICN | [5 empty] | 2187200
[20 empty] | VRE 4 ICN | [5 empty] | 2187200-2
| 3). Email all completed checklists to: @Health MR Upgrades or click: |  |  |  | mailto:@Health MR Upgrades & your regional sales/Mkt team | [20 empty] | 2187200-3
2187200-4
| Note: There is a different form for a 1.5T system. | [24 empty] | 2187200-5
SECTION 1:
9x - 14x
| GENERAL INFORMATION | [6 empty] | 15
16
| Site Name: | [6 empty] | 16.0-V02
HD23
System ID:
Date Checklist Completed:
Field Engineer Name:
FE SSO Number:
GON# (if known):
SECTION 2:
CURRENT SYSTEM INFORMATION
(Select one option from each drop-down menu or enter text when required.)
Current System Type:
Current Software Type:
Was the current system installed as a new
Forward Production or as an Upgrade?
For 8-16 channel systems: If LPCA contains legacy, green bendix connectors used for head and multi-coils, it's an UPGRADE.
| If at HD select |  | Sales - If cabinet number is 2357500 or 2357500-2, add M3335PM.
| Syscab Part Number: |  | If 5118496-5, -6, -7 or any other cabinet number, add M3335PS.
Magnet Type:
Information found on magnet rating plate.
System Type:
Site Type:
Current Gradient Driver:
Information found on rating plate on the gradient cabinet.
Is Gauss field at the
Pen Panel less than 200?
If Yes, continue.
If No, Pen Panel must be moved to a serviceable location in a <200G field.
HDx upgrades cannot occur unless Pen Panel is in a <200G field.
Current Channels:
Does this site have a Depopulated 16CH switch?
Review contents of HART Log for presence of Depopulated 16CH switch.
System has MNS Option?
Is the Mux board
| model number |  |  | Sales - If Mux board is 2274188 or 2274188-2
| 2274188 or 2274188-2? |  |  | add M3335BZ.
The Mux Board is in the RRF Chassis located in the System Cabinet. To locate the model number, open the rear of the
System Cabinet and look for the board in Slot 4 and 5 or Slot 13 and 14.
Have you performed "saveINFO"
Performing saveINFO is mandatory and important.
CADstream version available at site?
CADstream Server versions 3x, 4.0, 4.1 or 4.1.3 are NOT supported with HD23 software (SW).
For continued use CADstream Server with HD23 SW, purchase upgrade through M7000JL, M7000GJ or M7000JJ.
NOTE: For MNS (LX and EXCITE), their MNS features will no longer work after the upgrade
SECTION 3:
COMPLETE THIS SECTION ONLY IF CURRENT GRADIENT DRIVER IS ACGD
(If the current Gradient Driver is not ACGD, skip Section 3.)
| Is the Power Supply model number 2334862? |  |  |  | Sales - If No, order M3000GR.
SGA Power Supply Model Number:
Information in this box is not required for
systems at Excite or higher:
SGA Date Codes
Gradient Module #1:
Gradient Module #2:
Sales - If date is between
| Gradient Module #3: |  |  |  | April 1, 2002 and September 30, 2002:
Order Qty 1-3 M3090ND.
Record the date for all three components.
(Date is found on the rating plate shown below.)
SECTION 4:
COMPLETE THIS SECTION IF CURRENT SYSTEM IS Excite / HD / HDx / HDxt
Below mentioned Hardware upgrade is only applicable if sytem is upgrading to HDxt "Optima edition HD23"
| Does the ASC/CAM chassis on the system |  |  |  | If No,
| have a ASC NB I/F Board (P/N: 2391146)? |  |  |  | Order M3341AP x qty 1 for systems without MNS Option
Order M3341AR x qty 1 for systems with 8KW AMT MNS
Order M3341AS x qty 1 for systems with 4KW CPC MNS
| Select MNS Type on the system |  |  |  | Answer Required for question above
SECTION 5:
COMPLETE THE FOLLOWING SECTION IF CURRENT SYSTEM IS HDxt
Look at the Gradient Amplifier control boards and record the part number (P/N) and revision (Rev) number for each:
|  |  |  | Part Number |  |  | Rev Number
X - Gradient Amplifier control board P/N and Rev:
|  |  |  | Part Number |  |  | Rev Number
Y - Gradient Amplifier control board P/N and Rev:
|  |  |  | Part Number |  |  | Rev Number
Z - Gradient Amplifier control board P/N and Rev:
| If any of the Gradient Amplifier control board P/N is 2396050 with Rev 4 or lower | [5 empty] | Order M3340GZ. x qty 1
Below mentioned Hardware upgrade is only applicable if sytem is upgrading to HDxt "Optima edition HD23"
Select "Yes" or "No" for each selection below:
| MGD Slot 12 or CAM Slot 9 Front is SCP 3? |  |  |  | If No, order M3000PJ
| Host Computer is Z400? |  |  |  | If No, order M3341AH
RFS or System Cabinet Component Upgrade
| Are existing ICNs Part# 5127452 or 5127452-3? |  |  |  | If Yes, order M3341JD for 8Ch or 16Ch System
M3341JA for 32Ch system
| Quantity of installed ICNs at site |  |  |  | Answer required for question below
| Part Number of installed ICNs |  |  |  | Is ICN Part# 5127452?
Yes, add M3340HY - DC ICN Harvest Box-HDxt
If Qty of Installed ICNs = 4, add M3340HY x qty 4
If Qty of Installed ICNs = 2, add M3340HY x qty 2
If Qty of Installed ICNs = 1, add M3340HY x qty 1
Is ICN Part# 5127452-3?
Yes, add M3340HZ - 4100 ICN HarvestBoxHDxtDV
If Qty of Installed ICNs = 2, add M3340HZ x qty 2
If Qty of Installed ICNs = 1, add M3340HZ x qty 1