10 – TROUBLESHOOTING MULTICOIL TR BIAS FAULTS 10-1 General Theory Overview

TR bias, supplied through a dedicated line from the SSM to the Multicoil Select Switch and the Preamp Protect Module, is provided to switch the PIN diode switching circuitry for the individual coils inside the multicoil (phased-array coil) during a multicoil scan.  –5V is applied to the selected individual coils during the receive portion of the scan to reverse bias the multicoil PIN switch circuitry and thus tune the multicoil.  +5V is applied during the transmit portion of the scan to forward bias the PIN switch circuitry and thus detune the multicoil.  This switching also serves a dual purpose to connect or isolate coil outputs from various receive lines.  Also, forward bias (+5V) is also applied during the receive portion of some scans to detune certain multicoils that are not being used so as to prevent them from collecting noise and spurious signal that could degrade the combined image.  Biasing the coil in this way and then also switching the coil output from the receive line is sometimes referred to as switching the coil “hard” off.  This is a very effective way to insure that the coil output is isolated from the receive line.  A less effective way to minimize unwanted receive signal from a coil is to switch the coil “soft” off by leaving the coil reversed biased (tuned) but yet switched from the receive line.

10-2 Isolating Multicoil TR Driver Faults

1. Setup for a scan using the CTLTOP coil (even if site does not have this coil) and then [Manual Prescan] [Done].

2. Refer to the CTLTOP Multicoil diagram on the 8.X Service CD-ROM under Coils, Multicoil Control Diagrams, “1.5T Teledyne, USAI As Well As 1.0T CTL Coils”.   This is the first coil on the list.  This diagram can be used for 1.0T also.

Note that on the CTLTOP diagram, with the exception of the J20 MC SIGNALS connector supplying the multicoil driver signals on the SSM, all the pin numbers are furnished with the connectors.

10-2 Isolating Multicoil TR Driver Faults (Continued)

3. Use a Digital Multimeter (DMM) to measure for voltages between the designated Multicoil Driver pins and the common pins at J20 MC SIGNALS on the rear of the SSM.  See Table 10-1 below for a listing of all the Multicoil Driver pin assignments and expected voltages for CTLTOP coil selection.  If necessary, trace these through the chain all the way to the Preamp Protect Module in the Low Profile Carriage Assembly.

TABLE 10-1 J20 MC SIGNALS PIN ASSIGNMENTS AND EXPECTED VOLTAGES FOR CTLTOP COIL SELECTION

MULTICOIL DRIVER		PIN 		 Measured Voltage     	COMMON PINS

1				8			-5VDC		10,11,12,13,14,15

2				6			-5VDC		10,11,12,13,14,15

3				3			-5VDC		10,11,12,13,14,15

4				1			+5VDC		10,11,12,13,14,15

4. If the problem is missing multicoil drive voltage from the SSM then replace the CPD Board.

5. When finished reconnect any disconnected cables and then, if necessary, return to the flowchart.

102