Coldhead shield temp high but everything else is fine…..compressor may be the issue.

Magnet pressure going up check compressor charge.

Coldhead and recondenser should have no more than .3K difference.

No duty cycles may be a indication of coldhead is dead.