HFO troubleshooting tips

6- SHIMCONFIG DETAILS

The ShimConfig.cfg file is located in /usr/ng/w/config and contains the following information

for each coil:

• Harmonic naming order  • Current calculation; on or off • Existing coil current value

• Perturbation current value • Maximum power supply current

• Power supply or gradient coil name • Harmonic name

The vi or jot editor is used to edit this file.

The configuration file contains information for all the magnets. Illustration 6-1 highlights a

portion of the file that pertains to the OpenSpeed shim information.

To Communicate with the Z2 Power Supply

- Open a C-shell

- Type "cu z2supply" to connect.

- To read the software revision, type "gv"

- To reset the supply, type "re" (if software revision is 4.07 or higher)

- To read the current to supply, type "gc z5"  (8192 counts/1 amp, and 32767 is max)

- To set current, type "sc z5 (value in counts)"

- To exit, type "~."