Start Ghosting Troubleshooting

Freq. Encoding Problem

?

Smeared Phase Problem

?

Discrete Phase Problem

?

All Other Scans

?

EPI

Scan

?

Go to Sheet XX

Go to Sheet 7

Go to Sheet 6

Go to Sheet 4

Go to Sheet 2

## NOTE

Items with shadows are linked to other areas in or out of the chart

Page 1

# 1- Ghosting Troubleshooting -Draft

# 2 - Ghosting Troubleshooting - Draft

Page 2

Use the same protocol, change nothing except replacing the patient with a phantom

Possible protocol issue.  Explore with Apps or call support

Is the problem repeatable on a phantom?

Eliminate the graphical protocol (oblique) to have a purely axial scan.  Rescan.

Problem solved?

Yes

Yes

Was there a problem with the pre-scan?

(Patient size, movement physiological, breathing)

No

No

Inform customer

Yes

Done

No

Go to Page 3.

No

Change the offset to remain in FOV

Yes

Is the offset not in the Field of View?

Problem Solved?

Yes

No

Isolate the problem by using a scaled-down version of what the customer did.

# 3 - Ghosting Troubleshooting - Draft

Page 3

Problem still there?

Eliminate the effect of the coil

Scan done on Body Coil

Scan done on Head Coil

Try the same protocol on the Head Coil

Reproduced?

Try the same protocol on the Body Coil

Yes

Perform remaining tests using the head coil.

No

No

Problem with Head Coil (???????)

Problem with Body Coil (???????)

Go back to Body Coil configuration

Problem starts with multi-slice.  Keep only one slice in the scan.  Make this slice at the isocenter.

???????

From Page 2 or

Page 4

???????

Are failures symmetric about isocenter?

Not likely Gradient Related

No

Yes

Proceed to Gradient Chain Isolation Procedure,

Page XX (???????)

Both Head & Body Coils Fail

Body Coil Fails

Head Coil Fails

Eliminate effect of coil (Page 3)

???????

???????

???????

Are the failures plane dependent?

Both FSE & FGRE Fail

FGRE Fails

FSE Fails

Yes

No

See other possible reasons for failures on Page 1

Fails?

Run System Performance Test

Page 4

# 4 - Ghosting Troubleshooting - Draft

Other Questions for Ghosting…

Insert Gradient Chain Isolation Procedure here!

Based on these direct to Gradient Chain Isolation on RST or HSS

Shadow Signatures

- Gray Comp

- RF

- Shim Gradient

- Spike Noise

- Vibration

- Phantom

- Exciter Moving Metal

Questions:

- Whole / Zoom

- Head / Body

- FSE

- FGRE

- Intermittent failure or failure moving

- Symmetrical about isocenter

- Fail at isocenter and off isocenter?

- Test mode failing:

  - Echo shift

  - Phase

  - Magnitude

Continuation from Page 4

Page 5

# 5 - Ghosting Troubleshooting - Draft

No

Is the Null  tilted?

See “Discrete, No Null”, Page 8

Yes

No

Are the ghosts at the image’s edge?

See “Discrete Edges”, Page 8

See “Null, Tilted”, Page 10

Is there a dark area, or null, in the image?

Is the Null in the center of the image?

See “Null, Offset”, Page 10

See “Null, Center”, Page 11

Is the Null in offset to the left or right of the image?

Can you count the number of ghosts?)  (I.E. Is there a discrete number of ghosts?)

See “Phase: Smeared” Page 7

No

Yes

Page 6

No

No

No

Yes

Yes

Yes

Yes

# 6 - Ghosting Troubleshooting - Draft

No

No

Yes

Problem Resolved?

Does intensity of ghost increase with distance from isocenter

?

See other possible reasons for failures on Page 1

- RF Amplitude   or Phase

Yes

No

- Bo short time constant?

- Phase Correction?

Is there a Null in the smeared ghost?

No

Is the smearing along the edge on the phase axis?

Done

Yes

Yes

No

Proceed to Gradient Chain Isolation Procedure,

Page XX (???????)

Problem Resolved?

Could be Noisy Read Out Gradient.  Check:

- Incoming line voltage

- Ripple

“Phase: Smeared” From Page 6

Page 7

Yes

# 7 - Ghosting Troubleshooting - Draft

Call Support

Problem Solved?

No

Done

Problem Solved?

Check Eddy Currents

- Is the Bo time constant too short?

- Verify the Phase Correction

Yes

Yes

???????

Is scan EPI?

Is scan FSE?

“Discrete, Edges” From Page 6

Yes

Possible protocol issue.  Explore with Apps or call support

Done

- Is the Bo time constant too short?

- Verify the Phase Correction

Problem Solved?

“Discrete, No Null”, From Page 6

Page 8

No

No

Yes

Yes

No

# 8 - Ghosting Troubleshooting - Draft

# 9 - Ghosting Troubleshooting - Draft

Yes

No

Page 9

Phase Correction

Change the state of Phase Correction by either checking or un-checking the button for Phase Correction

Line Voltage / Frequency

Run Scan

Problem Solved?

Done

Yes

Return to Flowchart:

- Null, Center

- Null, Offset

- Discrete, Edges

- Discrete, No Null

- Smeared, Edges

Leave Phase Correction turned off

No

Run an un-gated scan (i.e. ?????)

Check the line frequency and voltage of incoming line power

Appropriate voltage and frequency?

Make necessary adjustments

Return to Flowchart:

- Null, Center

- Discrete, No Null

- Phase, Smeared

Check Cross- term Eddy Currents

No

Yes

Problem Solved?

Contact Support

Done

Problem Solved?

Yes

No

Bo dither time constant too short?

“Null, Tilted”

From Page 6

No

Yes

Problem Solved?

Contact Support

Done

Problem Solved?

Verify Phase Correction Constant

Check Eddy Currents

“Null, Offset”

From Page 6

Page 10

Yes

No

# 10 - Ghosting Troubleshooting - Draft

Contact Support

Done

Problem Solved?

Check each of the following.  Retest in between scans to see if problem has been resolved:

Linear Phase Correction

Coil Loading/Tuning

Check each of the following.  Retest in between scans to see if problem has been resolved:

Line Frequency

Cold Head

Vibration

Gradient Amps

Check each of the following.  Retest in between scans to see if problem has been resolved:

Eddy Current

Vibration

Phase Correction

EPI

MEMP/GRE

Periodic FSE

What type of scan was being performed?

“Null, Center”

From Page 6

Page 11

Yes

No

# 11 - Ghosting Troubleshooting - Draft

No

Done

Repair the failures noted in HSS tests

Problem Solved?

No

Run HHS Testing

- Open Service Desktop

- Click on “Image Quality”

- Click on “HSS Tool”

- Click on “Start here to start this tool”

Continue with “Null, Center” Trouble-shooting

HSS Failures?

Yes

No

Continue with “Null, Center” Trouble-shooting

Vibration Related

Problem Possibly Vibration Related (see below)

Run Scan

Problem Solved?

Turn off the Cold Head

Cold Head

Page 12

Yes

Yes

# 12 - Ghosting Troubleshooting - Draft

Yes

No

No

Yes

Done

Problem Solved?

No

Check Shorts in Grafidy 3

Yes

Fails?

Run Grad Long option on SPT Menu

Done

Problem Solved?

Full Grafidy 3  (EPI Cals, Check shim probe cals)

No

Yes

Call Support

Eddy Currents in Spec

Run Grafidy 3 to see if Eddy Currents are in Spec.

Verify the problem is Eddy Currents

Eddy Currents

# 13 -Ghosting Troubleshooting - Draft

Page 13

???

Don’t know where to link

???

Page 14

# Ghosting Troubleshooting ? - Draft

?

# 4a) Ghosting Troubleshooting - Draft

No

Page 4(a)

FSE Scan?

Yes

There is a certain amount of ghosting expected with FSE scans.  If your images look similar to those below, note that this is normal and probably not much can be done to improve the quality of the images.

Image

Breathing

Image

Vessel, etc.

Image

Eyelid Flutter

Rule out patient movement

Seeing Patient Motion?

Protocol Problem

Run SPT Testing Page XX

Did the customer follow correct guidelines for phase/frequency swapping?

Image

FSE Wrist

Image

FSE Knee

RF Chain Troubleshooting (Page XX)

Ghosts in all 3 image axes

Ghosts in one image axis

Run Grad Diagnostics

Gradient Troubleshooting (Page XX)

Where do ghosts appear?

Set up simple experiment:

- Head DQA

- Spin Echo

- 1.5 mm slice thickness

- 24 FOV

- Scan in pure axes for X, Y, Z

# 5a) Ghosting Troubleshooting - Draft

Page 5(a)

Problem with the Gradient.  Switch gradient.

Suspect Gradient Drive.

Does problem follow the switched axis?

Swap the gradient cable connecting to the axis in which the ghosted image appears with the gradient cable of a good axis (axis producing a clean image).  Re-run scan.

Gradient Troubleshooting

Page 6(a)

# 6a) Ghosting Troubleshooting - Draft