## Slide 1: MRI BASIC SYSTEM OVERVIEW
2014 Leadership Meeting

## Slide 2: Objectives:
To develop a basic understanding of the major components of an MRI system
To understand how a magnet is made and works

## Slide 3
Resistive
Types of Magnets
Magnets that are created by connecting it to a current source.
They use electricity to generate the magnetic field.
Very low fields 0.2-0.3T.
Low power and site requirements .
Can easily be de-energized.

## Slide 4: Types of Magnets
Permanent
Magnets made out of magnetic permanent materials (rare earth).
The major advantage is the low cost since magnet is always present and does not need cryogens or electricity.
Disadvantages include, weight, transport, low magnetic fields and disposing of the magnets.
Usually 0.2-0.45T (Low field)

## Slide 5: Types of Magnets
Electromagnets made from coils of superconductive wiring.
The coil needs to be immersed in cryogens and cooled to cryogenic temperatures in order to create the magnetic field by inducing a current and keeping the resistance at zero.
Very high fields 0.5T-20T.
High cost due to cryogens
Safety risk derived from high fields, cryogens and quenches.
Superconducting

## Slide 6
Normal pressure ≈ 4 psi (± .1 psi)
Liquid Helium
Coldhead
Liquid He
Gas He
Vacuum
-452 F°
4K
Image Reconstruction
Patient
Exam
Date
1/1/2014
Chest
Smith, P.
Reference
Sequence
Slice mm 2
Magnetic Resonance Imaging System
Energy Source
RF Transmitter
Data flow
Energy Detection
RF Receive Coil
Image Display
Subject to be Imaged
Energy Source to be Detected
Energy Detection Device
Image Reconstruction Mechanism
Image Display Device

## Slide 7: Magnet Vessel

## Slide 8: Tesla
1 Tesla = 10,000 Gauss
Earth magnetism 0.5 gauss
Most magnets are 30,000 to 60,000 stronger than the earth magnetic field
Most magnets 1.5T, a few 3T some old 1T
Many 7T or higher are for clinical research around the world right now. Mostly universities and research hospitals

## Slide 9
Magnet
RF
Gradients
Systems Operations
Computing
Energy Source
Spatial Encoding
System Functions
Coldhead
Compressor
Lines
Chiller
RF Power Supply
RF Amp
Gradient Power Supply
Gradient Amps (X,Y,Z)
System Operation
Applications
Reconstruction
Communications
Image Display
PACS
Movements
System Interface & Connectivity
Safety
Basic Functional Components of any Magnetic Resonant Imaging System
System
Imaging
Energy Detector
Surface Coil
Body Coil
T/R Coil

## Slide 10
Gradient Amplifier
X
Gradient Amplifier
Y
Gradient Amplifier
Z
Gradient Power Supply
RF Shield/Faraday Cage/Screen Room/Enclosure/Scan Room
Patient Table
Magnet
Gradient Coils
Magnet
RF Coil
RF Coil
Patient
Exam
Date
1/1/2014
Chest
Smith, P.
Reference
Sequence
Slice mm 2
Penetration Panel
RF Power Supply & Amplifier
Gradient Coils
Host
Imaging
Systems
Cabinet
Vent Stack
Surface Coil

## Slide 11: How do we make a magnet?
During ramping, a heater circuit is turned on causing a finite resistance to develop in resistor R. Once current begins to flow, the heater is turned off, and the resistance of R falls to zero, allowing the current to continue indefinitely through the superconducting loop in "persistence mode".
Ramping the magnet - Questions and Answers ​in MRI (mri-q.com)