[Tech Tip] GE MR Gather Load Data

Gather Load Data is a powerful troubleshooting tool used for rapid diagnosis of multicoil bias and dynamic disable faults.  GLD takes about 1 minute to run and doesn’t require any phantoms or scanning.  Each of the troubleshooting examples below take less than 15 minutes and will pinpoint your exact failed part.

> Service Browser

> Diagnostics

> Peripherals

> Driver Module Diagnostics

> Gather Load Data

To start a test, first change the Default output file name to something you will remember.  I typically use my initials, run number, and a brief description of what I’m testing (ex. kp1_ctl).  Leave all other settings at their default values.

Multicoil Error Troubleshooting:

- Identify the coil causing the error in the error log.

- Plug in the bad coil.

- Change the Default filename (1) and run GLD.

- Open a shell window.

- Open the results file by typing - gedit /usr/g/service/log/filename1_mc.txt

  - Look for the channel number the error log indicated has an issue and verify it’s numbers are different from the other good channels.

- Remove the bad coil and plug in another multichannel coil.

- Change the Default filename (2) and run GLD.

- Open the results file by typing - gedit /usr/g/service/log/filename2_mc.txt

  - If the bad channel from the first GLD run does not have different values from the good channels, order the coil from step 2.

  - If the bad channel from the first GLD run is also bad this time, the problem is in the system.  Continue.

- Copnnect the coil from step 2 and swap cables on J5 and J10 on the back of the Driver Module.

- Change the Default filename (3) and run GLD.

- Open the results file by typing - gedit /usr/g/service/log/filename3_mc.txt

  - If the bad channel did not move, order Driver Module.

  - If the bad channel moved, call TSM for support.  The problem is between the driver module and coil.

*** RESTORE ALL CABLES BACK TO ORIGINAL POSITIONS ***

Dynamic Disable Troubleshooting:

- Unplug any coils from the system.

- Identify the DD BLD (front/back, 1/2/3/4) that has an error and if it is a short or open.

- Change the Default filename (1) and run GLD.

- Open a shell window.

- Open the results file by typing - gedit /usr/g/service/log/filename1_dd.txt

  - Find DD line noted in error log and verify it has the bad value noted in the error.

- Swap the cables on BD1 with BD4, and BD2 with BD3 at the back of the Driver Module.

- Change the Default filename (2) and run GLD.

- Open the results file by typing - gedit /usr/g/service/log/filename2_dd.txt

  - If the bad value did not move, order a Driver Module.

  - If the bad value moved, continue.

- Restore the 4 cables on the back of the Driver Module to original positions.

- On the equipment room side of the pen panel swap J92 with J96, and J93 with J95.

- Change the Default filename (3) and run GLD.

- Open the results file by typing - gedit /usr/g/service/log/filename3_dd.txt

  - If the bad value did not move, order cable from Driver Module to Pen Panel.

  - If the bad value moved, continue.

- Restore the 4 cables on the equipment room side of Pen Panel to original positions.

- On the magnet room side of the Pen Panel swap J72 with J76, and J73 with J75.

- Change the Default filename (4) and run GLD.

- Open the results file by typing - gedit /usr/g/service/log/filename4_dd.txt

  - If the bad value did not move, order a Pen Panel DD Filter.

  - If the bad value moved, continue.

- Restore the 4 cables on the magnet room side of the Pen Panel to original positions.

- At the magnet swap the two front DD lines with each other AND the two rear DD lines with each other.

- Change the Default filename (5) and run GLD.

- Open the results file by typing - gedit /usr/g/service/log/filename5_dd.txt

  - If the bad value did not move, order cable from pen panel to magnet.

  - If the bad value moved, call a TSM for support.  Issue is in the body RF coil.

*** RESTORE ALL CABLES BACK TO ORIGINAL POSITIONS ***