# CT System Inspection for Toshiba:

System Type:	Aquilion 16

Inspection Date	11/19/2015

Engineer:		Mike Jackson

Location:              	2722 Walnut Ave

Suit or Room

Tustin, CA  92780

Contact Name:    	JOLIE CECERE

Contact Number:  	949-244-5151

Contract Type:  	ex: full contract, Time and Materials

Site Hours	 	8AM – 5PM

System Serial Number:			LUA10Y2160

VRDU or PDU Type:			Teal PCDU

Reconstruction Configuration:	STANDARD

Injector Model:				N/A

Additional Hardware:			N/A

Are There Operator Manuals?		Yes

Are There Service Manuals?		Yes

Operating System:			Windows XP Pro

Software Base:          			V3.38ER002

Software Apps Level:			Software Apps

Software Options Scan Side:		CARDIAC, DIM_EXP, D_PRT_BW, D_PRT_COLOR, MULTIVIEW, MWM_SCU, PGP, PRINT_SCU, QDSCAN, REALEC, REAL_PREP, REAL_VIEW, STRG_SCU

Software Options Display Side: 	ALATO_3DVIEWER, ALATO_MULTISLICECT, ALATO_STANDARD, DCM_CDR, D_PRT_BW, D_PRT_COLOR, MULTIVIEW, PGP, PRINT_SCU, QDSCAN, STRG_SCU, TIFF, TSB_DVD, WINPRT

Are There Copies of the Base, Apps and Options On Site? No

What is missing?

Is There a Backup Disk? Yes  Most Recent Date?  12/3/2014

X-ray tube model 		   	CXB-750D/5A

X-ray tube serial	            	82735-V0

X-ray Tube Manufacture Date	10/1/2010

Tube Installation Date		11/3/2010

Tube Slice Count			99150

Gantry Software Slice Count	99150

Gantry Hardware Slice Count	118235

Generator Type			SPELLMAN

Heat Exchanger Hours 		21089

Data Transfer System		MUDAT

Does the Site Have the Cart and Service Jigs?

NO

Does the Site Have All 7 Toshiba Phantoms?

NO

Are There Any Recent Issues With the System?

NO

# System Performance Check

Note any irregularities or issues at the bottom of the form.

# Bootup

Turn the system on from a powered off state.  Both the Display and the Scan side shoot boot up in less than 10 minutes

# Basic Operation

- Test all buttons and functionality on the gantry control panels.  (tilt, laser lights, couch movement)  Listen and feel for any abnormality in the couch vertical and horizontal movement and the gantry tilt movement.

- Test the foot pedals functionality

# Image Quality

- Toshiba phantom test – per BCTech PM sheet

- Water phantom 4x4 tests – per BCTech PM sheet

- Uniformity test – per BCTech PM sheet

- Low contrast – per BCTech PM sheet

- High contrast – per BCTech PM sheet

- Artifact check – after each scan look for streaks, rings or artifacts - these show how to perform the test at 0.5mm slices; perform each of the four tests at 1mm, 2mm and 4mm thickness altering the number of slices acquired as necessary.  An example on a 32 slice would be: 0.5mmx32, 1mmx32, 2mmx16, 4mmx8

  - LL phantom, S&V, LL FOV, 120kv, 300ma, 1 sec, 0.5mm slices, most slices allowed, no stack, FC 70

  - LL phantom, Helical, LL FOV, 120kv, 300ma, 1 sec, 0.5mm slices, most slices allowed, no stack, FC 70 - look at the middle slices of the phantom only

  - S  phantom, S&V, S FOV, 120kv, 300ma, 1 sec, 0.5mm slices, most slices allowed, no stack, FC 70

  - S  phantom, helical, S FOV, 120kv, 300ma, 1 sec, 0.5mm slices, most slices allowed, no stack, FC 70 – look at the middle slices of the phantom only

# Reconstruction

- Use existing raw data to stack up 1000 or more images in the recon queue.  Check that images reconstruction is complete.  Pay attention to reconstruction times to see if the completion time is reasonable for that system.

- Use the RDDUtil.exe tool to check that all Raw Data Disks are healthy; should be option 9 or 10 for windows systems

- Run hwtest.exe – no errors should occur (On windows systems where the HWtest data is not installed on the H drive there may be errors saying that data could not be loaded.)

  - For unix systems run 10 iterations of “all tests”

  - For windows systems run 10 iterations of “all tests but #3” which should be option 10

# High Voltage

- Use the XC tool to view an x-ray waveform using the signal show command.  Use 120kv, 300ma, 0.5 seconds on the small and then the large filament.  Don’t forget to check the HV resistor at this time.

- Run all Ifs for 80kv and 135kv at large and small filament.  No station should require more than three shots; if it does run the station again after interpolation to see if the issue is resolved.  Check to see if any of the adjusted IF data is more than 5 units different than the previously saved data; if it is re-run Ifs and check again.

- Check the XC board error log for recent errors

- With the tube heat below 20% perform an exposure, open the gantry front cover, shut off gantry power using the two back gantry breakers (NFB1 and NFB2).  Listen to the tube rotor; it should be smooth, not noisy and should spin without power for an extended time.

# System and Gantry

- Check the GTS log for any errors that occurred today or in the elog for previous errors

- Visually inspect the covers for serious cracks or defects

- Remove the largest covers and inspect for serious dust or signs of neglect

- Check to see if there are any visible frayed cables; especially in gantry rotation section

- Visually check the slip rings and brush for cracks or significant wear

- Check the PDU or VRDU for signs of neglect or issues with the breakers

- With the power off to the gantry, spin the gantry by hand to check for excessive resistance or grinding spots

# System Shutdown

Perform a system shutdown.  The Scon and Dcon should shut down within 10 minutes

# Systems for Purchase Only

| Final image quality and functional test completed |  |
| --- | --- |
| All system defects have been documented below |  |
| Patient information has been removed from the system |  |
| System is cleaned and disinfected if necessary |  |
| Gantry has either shipping lock in place or is bolted from the side |  |
| Couch is ready for shipping – at max or min height |  |
| Software and manuals are collected |  |
| Phantoms are empty and collected |  |
| User accessories are collected |  |
| Power is removed from the system |  |
| Cables are removed with both ends labeled |  |
| System is unbolted from the floor |  |
| Exit path is ready for system removal |  |

# Report any issues below:

AC, NO SOFTWARE, ONLY TOS PHANTOM